Raleigh/Wake City-County

Bureau of Identification

Crime Laboratory Division

DWI Chemistry Unit

BLOOD CHEMISTRY TECHNICAL PROCEDURES

Table of Contents

1: Technical Procedure for Quality Assurance	3

2: Technical Procedure for Uncertainty of Measurement	10

3: Technical Procedure for General Laboratory Equipment	15

4: Technical Procedure for pH Meter	25

5: Technical Procedure for Use of the Hamilton Microlab 625 Diluter to Prepare Samples for Blood Alcohol Determination Using Headspace Gas Chromatography	29

6: Technical Procedure for Biotage TurboVap LV Evaporator	34

7: Technical Procedure for Determination of Alcohol and Acetone in Blood by Headspace Gas Chromatography	37

8: Technical Procedure for Enzyme Linked Immunosorbent Assay (ELISA) as a Drug Screen	52

9: Technical Procedure for Solid Phase Extraction of Basic Drugs for GC-MS Analysis	64

10: Technical Procedure for Solid Phase Extraction of Drugs for GC-MS Analysis	72

11: Technical Procedure for Supported Liquid Extraction (SLE) of THC, THC-OH and THC-COOH for LCMS-MS Analysis	81

12: Technical Procedure for Gas Chromatography/Mass Spectrometry (GC-MS)	89

13: Technical Procedure for DWI Blood Chemistry Analysis	107

14: Technical Procedure for Solid Phase Extraction of Benzodiazepines for GC-MS Analysis	118

15: Technical Procedure for Liquid Chromatography/Mass Spectrometry (LCMS-MS)	126

16: Technical Procedure for Supported Liquid Extraction (SLE) of Acidic, Basic, and Neutral Drugs for LCMS-MS Analysis………………………………………………………………………………………………………...……148

# 1: Technical Procedure for Quality Assurance

- Purpose / Scope – This procedure provides direction for the receipt and quality assurance of laboratory supplies, equipment, reagents, reference collections, reference standards and reference materials that affect casework in the DWI Blood Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification Crime Laboratory

- Definitions

  - Quality control check – Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

  - Performance verification – The initial confirmation of the reliability of a previously or externally validated method or instrument.

  - Commercial reagent - A purchased solvent or chemical.

  - Certified Reference Material (CRM) – Reference material characterized by a metrologically valid procedure for one or more specified properties, accompanied by a certificate that provides the value of the specified property, its associated uncertainty, and a statement of metrological traceability.

  - Critical reagent - Chemicals or reagents which critically affect the quality of tests which do not have their reliability verified as part of the quality control checks in a DWI Blood Chemistry Unit Technical Procedure.

  - Prepared reagent - Mixture of two or more reagents or a dilution.

  - Reference standard - Measurement standard designated for the calibration of other measurement standards (reference standards or equipment.)

  - Reference material - Material sufficiently homogeneous and stable with reference to specified properties, which has been established to be fit for its intended use in measurement or in examination of nominal properties.

  - Primary reference material - Any reference material obtained from a commercial source which has documentation issued by the manufacturer certifying its chemical composition or has documentation stating the manufacturer’s specifications for the material.  This material may be certified reference material if available and practicable.

2.10. Secondary reference material - Reference material from a non-commercial source or from a commercial source which does not have authenticating documentation from the manufacturer or is derived from reference material.

2.11. Authenticating documentation - A certificate of analysis provided by the manufacturer certifying chemical composition or a statement of the manufacturer’s specifications or any published spectral data from an informed treatise generally accepted in the field that identifies a chemical substance.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

- Procedure for Laboratory Supplies and Commercial Reagents

  - Update the DWI Blood Chemistry Unit Chemical Inventory log maintained in the DWI Chemistry Unit folder on the shared drive when a commercial reagent is received and when it is emptied / disposed.

  - Upon being opened, commercial reagent containers shall be marked as opened along with the  initials of the Blood Drug Chemist and the date.

    - When a commercial reagent is transferred to another container it shall be labeled with the following:

      - Identity

      - Supplier and lot number

      - Initials of the Blood Drug Chemist

      - Date

      - Expiration date, if applicable.

      - Any additional information as required by the CCBI Health and Safety Manual.

- Procedure for Prepared Reagents

  - Reagents may be prepared in any amount provided that the component ratios in the DWI Blood Chemistry Unit Technical Procedure are kept constant.

  - Labeling

    - Lot numbers for containers of prepared reagents shall be assigned using lot number designations as specified in the DWI Blood Chemistry Unit Technical Procedure.

    - Containers of prepared reagents shall be labeled with the following:

        - Identity of the reagent

        - Initials of preparer

        - Date of preparation

        - Lot number

        - Expiration date

        - QCC due date

        - Any additional information as required by the CCBI Health and Safety 	Manual.

    - Each new container of prepared reagent shall be documented in the reagent log with the following:

        - Identity of the reagent

        - Lot number

        - Reference to the DWI Blood Chemistry Unit Technical Procedure    followed for preparation

        - Initials of preparer

        - Date of preparation

        - Expiration date

        - QCC result and supplier and lot number of any reference material used

        - Component(s) and supplier and lot number

  - Storage

    - Reagents shall be stored in closed containers and, if applicable, as specified in the DWI Blood Chemistry Unit Technical Procedure used for preparation.

  - Expiration Dates

    - Internal standard solutions, calibration solutions and verification solutions shall have the expiration date be set as the earliest expiration of a reference material component unless otherwise specified in the DWI Blood Chemistry Unit Technical Procedure used for preparation.  Internal standard solutions, calibration solutions, and verification solutions that have expired may be retested annually to extend the expiration date by a year.  All other prepared reagents shall expire three years after preparation unless otherwise specified in the DWI Blood Chemistry Unit Technical Procedure used for preparation.

  - Quality Control Checks

    - Prepared reagents shall be quality control checked according to the DWI Blood Chemistry Unit Technical Procedure used for preparation prior to or concurrent with their initial use.

    - Prepared reagents, other than internal standard solutions, calibration solutions and verification solutions, with an expiry greater than six months must be have QCC(s) repeated every six months to ensure reagent reliability.

    - Document quality control checks in the reagent log with the following:

      - Date performed

      - Initials of Blood Drug Chemist

      - Reference material, supplier and lot number

      - QCC result

      - Due date for next QCC

- Procedure for Reference Materials

  - Reference materials used in the Blood Drug Chemistry Unit for the identification of controlled substances and in quality control checks are critical reagents.

  - Prior to initial use in laboratory examinations, new reference material will be analyzed utilizing the instrumental techniques for which the Reference Material will be used for analytical purposes and any other analytical techniques necessary to verify chemical composition.  Data produced from the initial evaluation may be used in casework if determined to be acceptable.

  - Reference material will be suitable for use if the analytical data is comparable to available authenticating documentation provided by the supplier and authenticating data from another source.  If authenticating documentation is unavailable from one of these sources, the Technical Leader may approve the reference material for use based on available authenticating documentation and a documented evaluation describing the reasoning why the analytical data appropriately corresponds to the chemical characteristics of the reference material.  Reference materials that do not have authenticating documentation from a source other than the supplier will be indicated in the Chemistry Reference Material Log.  Reference material that does not meet these requirements may not be used in laboratory examinations.  During the Chemical Substances Inventory Audit, the Technical Leader will attempt to locate any additional authenticating documentation that may have become available after the reference material approval process.

    - New reference material received from a supplier which bears a lot number which has been previously evaluated according to these requirements does not need to be re-evaluated for suitability of use.

  - All authenticating documentation and analytical data will be maintained on the CCBI shared drive (i.e., S: drive). The date and the analyst who acquired the analytical data will be documented in the instrument logbook.

    - Reference material that does not have an expiration date or has passed the manufacturer’s expiration date will be re-evaluated prior to use using the intended analytical technique to ensure its stability.  It is not required to repeat all instrumental techniques to ensure chemical stability.

  - Reference materials shall be maintained in a secure location and stored according to the manufacturer’s instructions, if applicable.

  - Reference materials approved for use in laboratory examinations shall be documented in the Chemistry Reference Material Log. The Chemistry Reference Material Log will be maintained on the CCBI shared drive (i.e., S: drive).  The Chemistry Reference Material Log will reflect the date of approval for use in laboratory examinations, the initials of the person approving use, and the instrumental techniques (i.e., Infra-red Spectrophotometry and Mass Spectrometry) for which a Reference Material has been determined to be suitable.

    - The Chemistry Reference Material Log will reflect the current inventory of Reference Materials.  Chemists will record the date, their initials, and the purpose on the Chemistry Reference Material Log when a Reference Material is removed for use.  Chemists will record the amount removed according to the units listed in the Chemistry Reference Material Log.

  - An audit of the DWI Blood Chemistry reference materials shall be conducted annually according to the CCBI Administrative Procedure for Annual Quality Audits.

- Procedure for Critical Reagents

  - Negative blood

    - Negative blood shall be received and labeled by the Blood Drug Chemist as directed for commercial reagents, refer to Section 4, and stored according to the manufacturers recommendations.  Negative Utak blood  container(s) may be thawed and divided into smaller containers for individual use.  Mark the smaller containers with the original lot number followed by a unique, sequential letter.  Once placed in the refrigerator for thawing, the individual container shall be marked with an expiry of one month. Negative sheep blood will be stored in the refrigerator and shall be assigned an expiration date of 1 year from the date of receipt.

      - In the event that the container is too small to be labeled as required, the material may be stored in a larger container or the required information may be recorded in a Critical Reagent logbook.  Blood Drug Chemists shall refer to the Critical Reagent logbook for any information not recorded on the container prior to using the material.

    - Prior to use, for each new lot of negative blood a single sample must be analyzed according to the DWI Blood Chemistry Unit Technical Procedure(s) for the intended analysis technique(s).

      - The negative blood must not contain any controlled substances or controlled substance metabolites.  The presence of any non-controlled substances identified, e.g., caffeine, shall be recorded.

    - The data for negative Blood found to be suitable for use shall be marked with the date approved and the initials of the Blood Drug Chemist.  All analysis data and authenticating documentation shall be maintained in the Critical Reagent folder on the shared drive.

  - Chemical derivatizing agents

    - Chemical derivatizing agents shall be received and labeled by the Blood Drug Chemist as directed for commercial reagents, refer to Section 4, and stored according to the manufacturer’s instructions, if applicable.  Once opened, the container shall be used and disposed on the same day.

      - In the event that the container is too small to be labeled as required, the material may be stored in a larger container.  Blood Drug Chemists shall refer to the Critical Reagent folder for any information not recorded on the container prior to using the material.

    - Prior to use, for each new lot of chemical derivatizing agent, morphine reference material must be derivatized and analyzed by GC-MS.

      - The derivatized morphine mass spectrum must be substantially the same as reference material.

    - Chemical derivatizing agents found to be suitable for use shall be marked with the date and the initials of the Blood drug chemist.  All analysis data and authenticating documentation shall be maintained in the Critical Reagent folder on the shared drive.

- Procedure for In-house Generated Reference Collections

  - Spectral and relative retention time reference collections generated within the Laboratory will be traceable to primary reference materials, if practicable, otherwise secondary reference materials may be used.  Data and authenticating documentation shall be maintained in the DWI Blood Chemistry Unit folder on the shared drive.

- Procedure for Reference Standards

  - Refer to the DWI Blood Chemistry Unit Technical Procedure for General Lab Equipment

- Safety

  - Refer to the CCBI Health and Safety Manual

- Records

  - Reagent log

  - Chemistry reference material log

- References

  - ASCLD/LAB Policy on Measurement Traceability, ASCLD/LAB, AL-PD-3057 Ver 1.3.

  - ASCLD/LAB Guidance on Measurement Traceability, ASCLD/LAB, AL-PD-3058 Ver 1.0.

  - Part IVA.6.2 – Assessment of Drug Reference Materials. Scientific Working Group for the Analysis of Seized Drugs (SWGDRUG) Recommendations. 7th ed.: June 9, 2016.

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
|  |  | Updated section 3, updated lines 6.5. and 6.6.1. |
| 1/16/15 | 2 | Update and additions to 7.1, 7.1.1.1, 7.1.2, 7.2.1, 7.2.1.1 and 8.2. |
| 1/26/18 | 3 | Updates to 4.1, 4.2, 4.3, 4.4, and 5.2.2 to be consistent with changes in Laboratory Administrative Procedures Manual. |
| 4/10/2018 | 4 | Updated safety manual name in 10. |
| 10/28/2019 | 5 | Updated 7.1.2 for change to SLE and updated 8.2 to remove THC and THC-COOH from RRT document. |
| 01/06/20 | 6 | 2. Definitions - added Certified Reference Material, removed purchasing documentation. 3. Abbreviations - removed abbreviations Removed any instructions which are already included in other laboratory procedures. 6. Reference Material - revised content 7. Procedure for Critical Reagents - revised labeling requirements for Negative blood and chemical derivatizing agents 11. Records - added Chemistry reference material log 12. References - revised references |
| 6/5/20 | 7 | 5.4 Expiration Dates - revised expiration date criteria 8. Procedure for in-house generated Reference Collections - removed 8.2 and 8.3 |
| 10/1/20 | 8 | 7.1. Negative Blood - added requirements for storage of negative sheep blood to 7.1.1. and analysis requirements in 7.1.2., and 7.1.3. to allow for electronic storage of data. 7.2. Chemical derivatizing agents - modified 7.2.1, 7.2.1.1. and 7.2.3. to allow for electronic storage of data. |

# 2: Technical Procedure for Uncertainty of Measurement

- Purpose / Scope -This procedure is utilized to estimate the uncertainty of measurement for test methods for which a numerical value is reported on a Laboratory Report in the DWI Blood Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification Crime Laboratory.

- Definitions

  - Uncertainty of measurement - a parameter associated with the result of a measurement that characterizes the dispersion of the values that could reasonably be attributed to the measurand.

  - Coverage probability (Level of confidence) - probability that the set of true quantity values of a measurand is contained within a specified coverage interval.

  - Coverage factor - numerical factor used as a multiplier of the combined uncertainty in order to obtain an expanded uncertainty.

- Abbreviations

  - UOM – uncertainty of measurement

  - CU – combined uncertainty

  - EU – expanded uncertainty

- Procedure

  - The DWI Blood Chemistry Technical Leader shall determine an estimation of the UOM for each test method for which a numerical value is reported on a laboratory report.  The specific measuring device or instrument used for a reported test result must be evaluated in the estimation of the UOM for that test method.

  - The estimation of the UOM shall be performed annually, at a minimum, or when a change in measurement conditions occurs that may have a significant effect on the UOM.

  - Laboratory environmental conditions shall be monitored and any additional effect on UOM shall be evaluated upon collection of data.  Refer to the DWI Blood Chemistry Unit Technical Procedure for General Laboratory Equipment.

  - Each test method requiring UOM shall be evaluated for contributions from sources of uncertainty, u.  The contributions shall be evaluated using Type A methods (by a statistical analysis of measured values obtained under defined measurement conditions such as repeatability and / or reproducibility, including measurement assurance data) and Type B methods (by other means of analysis of components from such things as instrument readability, calibration certificate reported uncertainty, etc.)

  - Evaluate the identified sources of uncertainty and combine them to obtain the combined uncertainty of measurement, CU,  using the formula

CU = √(u12 + u22 + u32…. )

where

CU = combined uncertainty

u1, u2, etc.  = individual identified sources of uncertainty

  - The combined uncertainty of measurement is an estimation of the uncertainty of measurement, UOM.  Individual sources of uncertainty that are not significant contributors may be excluded.

  - The expanded uncertainty, EU, shall be calculated to provide a minimum 99.73 % coverage probability (or approximately 99%) by multiplying the CU by the appropriate coverage factor, k.

  - Round the EU to two significant digits.  Do not perform rounding prior to this step.

    - When the digit next beyond the one to be retained is less than five, keep the retained figure unchanged.  For example:  2.541 becomes 2.5 to two significant figures.

    - When the digit next beyond the one to be retained is greater than five, increase the retained figure by one.  For example:  2.453 becomes 2.5 to two significant figures.

    - When the digit next beyond the one to be retained is exactly five, and the retained digit is even, leave it unchanged; conversely if the digit is odd, increase the retained figure by one (even/odd rounding).  Thus, 3.450 becomes 3.4 but 3.550 becomes 3.6 to two significant figures.

    - When two or more figures are to the right of the last figure to be retained, consider them as a group in rounding decisions.  Thus, in 2.4501, the group (501) is considered to be greater than 5 while for 2.5499, (499) is considered to be less than 5.

  - For blood alcohol and acetone determination, multiply the % EU by the average of the four measured values (gram / 100 ml), to obtain an EU expressed in the same units as the measurement result.

  - The reported EU shall contain at most two significant digits and be reported to the same level of significance as the measurement result.  Any requirements for the level of significance for reporting of the measurement result shall also determine the level of significance for reporting the EU.  The reported EU shall be rounded, refer to 4.8.1 – 4.8.4 for rounding.

    - For blood alcohol and acetone determination, the blood alcohol concentration is truncated to the hundredths place according to North Carolina General Statutes § 20-4.01.(1b).  The EU is not reported for the statutorily required blood alcohol concentration.

  - The EU shall be reported for each test method where a numerical value is reported on a laboratory report.  When numerical results are added to produce a combined result the respective EU’s shall also be added.

  - In addition to the statutorily required blood alcohol report the laboratory report shall identify the measured quantity value, y, along with the associated EU.  The report shall be reported as y ± EU, with the units of EU consistent with the units of y. For blood alcohol and acetone determination, round the EU determined in 4.9 to a maximum of  two significant figures and round the average of the measured values to the same level of significance (decimal places).  Refer to 4.8.1 – 4.8.4 for rounding.

    - Add the statement containing the values from 4.12 to the report.

      - Example:

(The average of the four measured alcohol concentrations is 0.1034 gram per 100 milliliters ± 0.0062 gram per 100 milliliters of whole blood at a coverage probability of 99.73%.)

(The average of the four measured alcohol concentrations is 0.374 gram per 100 milliliters ± 0.022 gram per 100 milliliters of whole blood at a coverage probability of 99.73%.)

(The average of the four measured alcohol concentrations is 0.0123 gram per 100 milliliters ± 0.0007 gram per 100 milliliters of whole blood at a coverage probability of 99.73%.)

  - The DWI Blood Chemistry Unit Technical Leader shall maintain records of the estimation of the uncertainty of measurement in the DWI Blood Chemistry Unit.  The records shall:

- Define the measurand

- State how traceability is established for the measurement

- State the equipment (measuring device(s)) used

- State all uncertainty components considered

- Identify all uncertainty components of significance and how they were evaluated

- Contain the data used to estimate repeatability and / or reproducibility

- Contain all calculations performed

- State the combined standard uncertainty, CU, the coverage factor, k, the coverage probability, C, and the resulting expanded uncertainty, EU

- The minimum due date for the review/recalculation of the measurement uncertainty, refer to 4.2.

- Calculations

  - CU = √(u12 + u22 + u32…. )

  - Average or mean, x = 1n i=1nxi

- Records

  - Measurement Uncertainty for Blood Alcohol and Acetone

  - Uncertainty of Measurement Budgets

- References

  - ASCLD/LAB Level 100A Traceability presentation, Copyright 2011; Heusser Neweigh, LLC & ASCLD/LAB.

  - ASCLD/LAB Level 100B Measurement Assurance presentation, Copyright 2011; Heusser Neweigh, LLC & ASCLD/LAB.

  - ASCLD/LAB Level 100C Measurement Uncertainty Concepts presentation, Copyright 2011; Heusser Neweigh, LLC & ASCLD/LAB.

  - ASCLD/LAB Level 200 Measurement Confidence for the Forensic Laboratory: Measurement Uncertainty in Drug Chemistry presentation, Copyright 2011; Heusser Neweigh, LLC & ASCLD/LAB.

  - ASCLD/LAB Level 200 Measurement Confidence for the Forensic Laboratory: Measurement Uncertainty in Toxicology Testing presentation, Copyright 2011; Heusser Neweigh, LLC & ASCLD/LAB.

  - Introduction to Measurement Uncertainty course, LeBeau, Marc A., Ph.D., 2009, RTI International

  - Introduction to Measurement Uncertainty – Practical Examples Part II course, LeBeau, Marc A., Ph.D., 2010, RTI International

  - Introduction to Measurement Uncertainty – Practical Examples Part III course, LeBeau, Marc A., Ph.D., 2010, RTI International

  - Evaluation of measurement data - Guide to the expression of uncertainty in measurement, JCGM 100:2008 GUM 1995 with minor corrections, First edition September 2008, JCGM 2008, Working Group 1 of the Joint Committee for Guides in Metrology (JCGM/WG 1)

  - ASCLD/LAB Policy on Measurement Uncertainty, ASCLD/LAB, AL-PD-3060 Ver 1.1.

  - ASCLD/LAB Guidance on the Estimation of Measurement Uncertainty – Overview, ASCLD/LAB, AL-PD-3061 Ver 1.0.

  - ASCLD/LAB Guidance on the Estimation of Measurement Uncertainty – ANNEX A, Details on the NIST 8-Step Process, ASCLD/LAB, AL-PD-3062 Ver 1.0.

  - ASCLD/LAB Guidance on the Estimation of Measurement Uncertainty – ANNEX D, Toxicology Testing Discipline Example – Concentration of Ethanol in an Ante-Mortem Blood Specimen, ASCLD/LAB, AL-PD-3065 Ver 1.0.

  - GLP 9 Good Laboratory Practice for Rounding Expanded Uncertainties and Calibration Values, National Institute of Standards and Technology (NIST), Gaithersburg, Maryland, February 2012.

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
|  |  | Updated section 3. |
| 12/16/13 | 2 | Compliance with new ASCLD/LAB policies |
| 4/14/14 | 3 | Add report statement for compliance with statutory limits |
| 10/1/14 | 4 | Removed rounding in 4.9 |
| 1/16/15 | 5 | Corrected typographical error in 4.8.2. Added clarification in 4.13 and example in 4.13.1.1. Added additional examples of statutory limits in 4.14. |
| 9/15/17 | 6 | Removed uncertainty statement for statutorily reported alcohol concentration values in sections 4.10.1 and 4.12. Added wording from 4.12 to 4.13. Updated wording in 4.12.1 Removed section 4.14 to be consistent with DBCTP13 |
| 1/26/18 | 7 | Added Section 6. Records |
| 10/28/19 | 8 | Added 4.10 to add blood drug EU criteria and example. |

# 3: Technical Procedure for General Laboratory Equipment

- Purpose / Scope - This procedure provides direction for the use of general laboratory equipment in the DWI Blood Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification.

- Definitions

  - Quality control check – Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

  - Performance verification – The initial confirmation of the reliability of a previously or externally validated method or instrument.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

  - NIST – National Institute of Standards and Technology

  - QCC – Quality control check

- Equipment, Materials

  - Equipment

    - Refrigerator

    - Freezer

    - Eppendorf Reference Mechanical Pipettes and corresponding pipette tips

    - Mettler XS1003S Balance

    - Class A volumetric glassware, pipettes and volumetric flasks

    - Thermometers, NIST traceable

    - Laboratory environmental monitor – barometer, hygrometer, thermometer, NIST traceable

    - Millipore Elix Advantage Water Purification System

  - Materials

    - Weighing vessels

    - Deionized water

    - Millipore Elix Advantage Progard Pack

    - Millipore Elix Advantage Reservoir Vent Filter

    - Millipore Elix Advantage UV 254 lamp

    - Millipore Elix Advantage chlorine tablet

  - Reference Standards

    - NIST traceable weights:  0.1 g, 2 g, 200 g

- Procedure for Refrigerators and Freezers

  - Monitor the temperature of each refrigerator used for storage of evidence, chemicals or reference materials each working day.

  - Record the temperature in degrees Celsius on the refrigerator or freezer temperature log.

  - The acceptable temperature for refrigerators is 40C (± 30C).  The acceptable temperature for freezers is -150C or below.

    - If the temperature is unacceptable adjust the temperature setting.

    - Notify the Blood Drug Chemistry Technical Leader for scheduling of service or replacement if the temperature cannot be corrected.

      - Mark improperly functioning evidence refrigerators “out of service,” relocate the evidence and record in the case record

  - Record all adjustments, maintenance and comments in the refrigerator or freezer maintenance log.

  - Maintain refrigerator and freezer logs in the DWI Blood Chemistry Unit.

- Procedure for Mechanical Pipettes

  - Mechanical pipettes shall be certified by the manufacturer and approved prior to use.

  - After one year of use the mechanical pipettes shall be serviced and calibrated annually by an approved vendor.  Upon completion of calibration a DWI Blood Chemistry Technical Leader will review and sign the calibration report.

  - Perform performance verification on all mechanical pipettes used for evidence handling, i.e., 1 ml and 0.25 ml pipettes, prior to their initial use.  Label the pipettes “Blood.”  Follow the procedure for a monthly quality control check, refer to 6.4.1.

  - Perform a monthly quality control check prior to the first use of each month on all “Blood” mechanical pipettes.

    - Monthly Quality Control Check

      - Inspect the pipette and tips for proper functioning.

        - If a pipette problem is observed, remove the pipet from service, label it “out of service,” and notify the DWI Blood Chemistry Technical Leader for scheduling of repair.

        - If pipette tip problems are observed, segregate the pipette tips and clearly label “not for use.”  Notify the DWI Blood Chemistry Technical Leader.

      - Allow the pipettes, pipette tips, and deionized water to equilibrate in the measurement location for a minimum of two hours at room temperature.

      - Place the pipette tip on the tip cone of the pipette.

      - Fill the weighing vessel with deionized water to a minimum height of 3 millimeters.

      - Fill the pipette tip and empty five times with the deionized water.

      - Discard the pipette tip and replace.

      - Pre-wet the tip one time.

      - Hold the pipette vertically.

      - Dip the pipette tip into the deionized water by a few millimeters.

      - Draw in the volume to be tested slowly and evenly and wait three seconds.

      - Pull the pipette tip slowly out of the liquid, wiping it on the vessel wall.

      - Rest the filled tip up against the wall of the weighing vessel at an angle.

      - Dispense the test liquid slowly until the first stop (measuring stroke).

      - Press the control button to the second stop (blow-out) and dispense the remaining liquid in the tip.

      - Hold down the control button and pull the tip up the vessel wall.

      - Let the control button slide back into position.

      - Determine and record the weight.

      - Repeat nine times.

      - Calculate the systematic error in percent.

        - Calculate the mean value, x, of the dispensed volume, using the correction factor, z = 1.0030.

x =z All measured valuesn

where n = number of measured values

        - Calculate the relative systematic error, es, in percent.

es=100x- xnominalxnominal

where xnominal = target volume

        - The relative systematic error will be truncated to one decimal place for evaluation. The acceptable range is ∓ 5 %.

      - Calculate the relative random error as coefficient of variation.

        - Calculate the absolute random error, s, using the correction factor, z = 1.0030.

s =xiz-x2n-1

        - Calculate the coefficient of variation, CV (relative random error).

CV (%)= 100sx

        - The coefficient of variation will be truncated to one decimal place for evaluation. The acceptable range is ∓ 5 %.

      - Record the relative systematic and random errors on the Pipette Quality Control Check form.

      - If an error is outside of the acceptable range, the QCC may be repeated.  If the problem persists, remove the pipette from service, label it “out of service,” and notify the DWI Blood Chemistry Technical Leader for scheduling of repair.

  - After completion of any service from an outside vendor perform the Monthly QCC prior to placing the pipettes back in service.

  - For each “Blood” mechanical pipette, perform a Daily Quality Control Check prior to use for casework. Dispense and weigh one sample of water. The water sample must weigh within +/- 5% of the expected weight of water to be delivered (1mL – 0.950-1.050g and 0.25mL – 0.238-0.262g). Record this weight on the Pipette Quality Control Check form. A Monthly QCC will serve as the Daily QCC for the first use of the month.

  - Prior to each use inspect the pipette and tips to ensure that the control button is moving properly, a visually adequate volume of liquid is aspirated and delivered, there is no dripping.

    - If a pipette problem is observed, remove the pipet from service, label it “out of service,” and notify the DWI Blood Chemistry Technical Leader for scheduling of repair.

    - If a pipette tip problem is observed, segregate the pipette tips and clearly label “not for use.”  Notify the DWI Blood Chemistry Technical Leader.

  - Pipette operation

    - Pre-wet pipette tips by aspirating and dispensing the liquid to be measured a few times before pipetting.  With the tip not in contact with the liquid, empty it completely on the inner wall of the tube (via blow-out).

    - Aspirating

      - Attach corresponding pipette tip to pipette.

      - Press down the control button to the first stop (measuring stroke).

      - Immerse the pipette tip vertically approximately 3 millimeters into the liquid.

      - Allow the control button to slide back slowly.

      - Pull the tip out of the liquid slowly.

      - To remove any droplets, dab with line-free wipe, ensuring that no liquid comes out of the tip.

    - Dispensing

      - Hold the tip at an angle against the inside wall of the tube.

      - Press down the control button slowly to the first stop (measuring stroke) and wait until the liquid stops flowing.

      - Press down the control button to the second stop (blow-out) until the liquid stops flowing.

      - Hold down the control button and pull the tip up the inner wall of the tube.

      - Eject tip by pressing the control button to the final stop.

  - Record maintenance and service in a log maintained in the DWI Blood Chemistry Unit.

  - Maintain manufacturer’s certificates, performance verification documentation, quality control check documentation and calibration certificates in the DWI Blood Chemistry Unit.

- Procedure for Balance

  - Balances shall be certified by the manufacturer and appropriate for the intended use.

  - Calibration for balances shall be performed on an annual basis by an approved vendor that meets the requirements specified by the accrediting body.

  - Monthly Quality Control Check

    - Perform a Monthly Quality Control Check prior to the first use of each month using the 200g, 2g, and 0.1g reference weights.

    - Ensure that the balance is clean and level.  If the balance is not level, adjust the leveling  screws until the air bubble is within the inner circle of the level indicator.

    - Press the internal adjustment key and allow the balance to complete the internal adjustment function.

    - Zero the empty balance and allow the reading to stabilize. Ensure that the reading is zero.

    - Place a reference standard weight on the center of the balance and allow the reading to stabilize. Record the displayed weight on the Balance QCC Log.  The acceptable range for each weight is listed below.

| Reference Standard Weight | Acceptable Range |
| --- | --- |
| 200 g | 199.998 - 200.002 g |
| 2 g | 1.998 - 2.002 g |
| 0.1 g | 0.098 g - 0.102 g |

    - If the results are outside of the acceptable range, correct any apparent problem utilizing the necessary steps to bring the balance back into compliance. Steps may include cleaning, leveling, zeroing and performing the internal adjustment.

      - If a problem was observed and corrected, record the correction activity and repeat the QCC.

      - If the repeated QCC remains outside of the acceptable range, label the balance and QCC Log “out of service” and notify the DWI Blood Chemistry Technical Leader.

  - Balance operation

    - Leave the balance powered on.

    - Perform a Daily Quality Control Check on the balance using the 2 gram reference standard weight prior to use for casework, refer to 7.3. A Monthly QCC will serve as the Daily QCC for the first use of the month.

- Procedure for Class A Volumetric Glassware

  - Volumetric glassware used in Blood Alcohol and Cannabinoid analyses and used to prepare internal standard solutions shall be Class A.

  - Inspect the glassware upon each use for damage and cleanliness.

    - Discard glassware that is damaged, e.g., chipped or cracked.

    - Discard glassware that cannot be thoroughly cleaned, e.g. discard pipets that do not drain cleanly, discard flasks that cannot be cleaned.

  - Maintain certificates and/or manufacturer specifications for volumetric glassware in the DWI Blood Chemistry Unit.

- Procedure for Laboratory Environmental Monitor

  - The building environmental controls generally provide adequate laboratory environmental conditions for analysis.  When environmental conditions are outside of acceptable levels they are apparent to a Blood Drug Chemist based on their sensory perception of the environment.  When a Blood Drug Chemist perceives an abnormally high or low temperature or abnormally high humidity level the Blood Drug Chemist shall:

- Notify General Services Administration that there is an environmental control problem and provide the applicable room number

- Notify the DWI Blood Chemistry Technical Leader and Crime Laboratory Quality Manager of the environmental control issue

- Measure the laboratory environmental conditions using a hygrometer and thermometer.

  - Record the humidity, temperature and air pressure on the environmental logs.

  - If the humidity exceeds 80% or the temperature is not in the range of 15.5 – 32.2 °C, stop all analyses until environmental conditions return to acceptable levels.

    - All instruments shall be subjected to any post shutdown checks as described in the DWI Blood Chemistry Unit Technical Procedures.

- Millipore Elix Advantage Water Purification System with E-POD Unit

  - Prior to dispensing deionized water ensure that the resistivity displayed on the E-POD unit is greater than 1.0 MOhm.cm.

  - To dispense water press down on the E-POD Unit plunger.  Push slightly for low flow.  Push down and hold for high flow.  Push completely down for continuous high flow and push down again to stop.

  - Maintenance

    - Replace the Progard Pack, Reservoir Vent Filter and UV 254 lamp when prompted by the System.  Refer to 14.9.

    - Clean the Inlet Strainer and RO Cartridge when prompted by the System.  Refer to 14.9.

    - Record maintenance in a logbook maintained in the DWI Blood Chemistry Unit.

  - Record System alarms in the logbook and notify the DWI Blood Chemistry Unit Technical Leader if service is required.

- Safety – Refer to the CCBI Health and Safety Manual

- Records

  - Blood Chemistry Balance log

  - Blood Chemistry Balance quality control check form

  - Refrigerator temperature log

  - Refrigerator maintenance log

  - Freezer temperature log

  - Freezer maintenance log

  - Environmental log

  - Pipette quality control check form

  - Pipette log

- References

  - Eppendorf Reference Operating Manual, 2001, Eppendorf AG, Hamburg.

  - Eppendorf Standard Operating Procedure for Pipettes, 2016, Eppendorf AG, Hamburg.

  - Instructions for Installation, Use & General Maintenance of Your Glass Door Merchandiser MAN-004-R, 2009, Fogel.

  - T° Sentry Model 140 Digital Alarm Module Operating Instructions, 2004, Hampshire Controls Corporation.

  - Kenmore Elite Refrigerator Use & Care Guide, 2009, Sears, Roebuck and Co. Hoffman Estates, IL.

  - Thermo Scientific General Purpose, Flammable Storage, and Explosion Proof Laboratory Refrigerators and Freezers Installation and Operation Manual, 2010, Thermo Fisher Scientific.

  - Operating Instructions for Excellence Precision Balances XS Models – Part 1, 05/2012, Mettler-Toledo, Greifensee, Switzerland.

  - Elix Advantage 3/5/10/15 System User Manual, Millipore Corp, France, FTPF11337-V1.0, 02/2010.

  - Operating Instructions for Excellence Balances XS Models – Part 2, 10/2010, Mettler-Toledo, Greifensee, Switzerland.

  - Drug Chemistry Unit Technical Procedure for Balances.

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
| 3/27/13 | 2 | Incorporated new balances. |
| 1/16/15 | 3 | Minor wording changes. Update record storage requirements and refrigerator temperature range. Remove reference to Millipore POD Pak. |
| 11/06/2015 | 4 | Update 10.1. |
| 4/10/2018 | 5 | Updated criteria for pipette acceptance in Monthly QCC in sections 6.4.1.20.3 and 6.4.1.21.3. Updated 6.5 to require for a QCC after service. Removed Dispenser Procedure from Section 9, Dispenser from 4.1.6 and log from 13.10. Updated Balance and QCC log form name in 13.1 and 13.2, respectively. Updated safety manual name in 12. Updated reference 14.2. Removed Dispenser reference from 14.3. |
| 5/8/2018 | 6 | Added requirement for pipette calibration reports to be signed. |
| 3/1/2019 | 7 | Revised 6.4.1 to remove requirement of specific temperatures and recording of temperature and pressure using the environmental monitor; added a constant correction factor of z=1.0030 based on historical data; added requirement to evaluate relative systematic error and relative random error based on truncated values; Added 6.6 requirement for daily quality control check prior to use; added acceptable daily usage check ranges; Condensed section 7 Procedure for Balance and added daily QCC requirement |

# 4: Technical Procedure for pH Meter

- Purpose / Scope - This procedure provides direction for the use of the pH meter in the DWI Blood Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification.

- Definitions

  - Quality control check – Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

  - Performance verification – The initial confirmation of the reliability of a previously or externally validated method or instrument.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

  - QCC – Quality control check

- Equipment, Materials, Reagents

  - Equipment

    - pH Meter, Fisher Scientific Accumet Basic AB15/15+, with double-junction combination pH electrode and ATC probe

  - Materials

    - Deionized water

    - Beaker

    - Wash bottle

  - Reagents

    - pH buffers, certified by manufacturer:  4.0, 7.0 and 10.0

    - Electrode storage solution, potassium hydrogen phthalate-potassium chloride

    - Electrode fill solution, saturated potassium chloride

- Standardization of pH Meter

  - Standardize the pH meter daily, when in use, with two buffer solutions with values that bracket the desired measuring range, e.g., pH 4.0 and pH 7.0.

  - If necessary, fill the electrode with the electrode fill solution.

  - Ensure that the filling hole of the electrode is open.

  - Press and release the mode key until the digital display indicates pH Mode.

  - Press the setup key twice and then press enter key to clear an existing standardization.

  - Immerse the rinsed electrode and temperature probe into a beaker containing a buffer solution and stir moderately.

  - Press the std key to access the Standardize mode. A selected buffer group is displayed briefly.

  - Wait for reading to stabilize.

  - Press the std key again to initiate standardization. The meter will automatically recognize the buffer and then the meter returns to the measure screen.

  - Repeat 5.6 -5.8 with a second buffer solution. When the meter accepts the second buffer, it will briefly display the percent slope associated with the electrode’s performance prior to returning to the measure mode. If the electrode is within the range of 90% - 102%, the “good electrode” message will appear. If the electrode is outside this range, the meter will display the “electrode error” message and will not return to the measure screen until the user presses the enter key.

    - If the electrode is outside of the “good electrode” range of 90% - 102% an “electrode error” message will appear.  Rinse and fill the electrode with electrode filling solution and repeat the standardization.

      - If the problem persists test the meter for correct operation.

      - Install the BNC (input) shorting cap.

      - Press mode to access the mV mode and record the mV reading.  If the meter reads 0 ± 1 mV the meter is measuring correctly, proceed to 5.10.1.4.  If the reading is outside the acceptable range, mark the pH meter log “out of service” and notify the DWI Blood Chemistry Technical Leader for scheduling of repair.  Record the activity and result in the pH meter log.

      - If the pH meter is functioning correctly replace the electrode, fill it with electrode filling solution, soak it for 2 – 4 hours in electrode storage solution and perform the standardization with the new electrode.  Record the activity in the pH meter log.

  - Record date, initials, buffers and lot numbers, slope and message displayed in the pH meter log.

  - Remove the electrode and temperature probe from the buffer solution and rinse with distilled water.

- Procedure

  - Prior to use daily standardize the pH meter, refer to section 5.

  - Ensure that the filling hole of the electrode is open.

  - Set the mode key to pH measure.

  - Rinse the electrode and temperature probe with deionized water and blot dry.

  - Immerse the electrode and temperature probe into a beaker containing the solution and stir moderately.

  - Wait for reading to stabilize and then record the reading.

  - Remove the electrode and thermometer from the solution.

  - Rinse the electrode and thermometer with deionized water before and blot dry.

  - Immerse the electrode in the electrode storage solution when the pH meter is not in use.

  - Record the date, initials, measured liquid identity and observed pH in the pH meter log.

  - Maintain the pH meter log and buffer certificates in a pH meter logbook in the DWI Blood Chemistry Unit near the pH meter.

- Safety

  - Refer to the CCBI Health and Safety Manual

- Records

  - pH meter log

- References

  - Accumet Basic AB15 User Manual, August 2006, Fisher Scientific.

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
|  |  | Updated Section 3, corrected numbering. |
| 4/10/2018 | 2 | Updated safety manual name in 7. |

# 5: Technical Procedure for Use of the Hamilton Microlab 625 Diluter to Prepare Samples for Blood Alcohol Determination Using Headspace Gas Chromatography

- Purpose / Scope - This procedure provides direction for the initial setup, maintenance, performance checks and usage of the Hamilton Diluter in the DWI Blood Drug Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification.

- Definitions

  - Performance verification - The initial confirmation of the reliability of a previously or externally validated method or instrument.

  - Quality control check (QCC) - Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

  - Reference Material – Material sufficiently homogeneous and stable with reference to specified properties, which has been established to be fit for its intended use in measurement or in examination of nominal properties.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

  - BAC – Blood alcohol content

  - QCC – Quality Control Check

- Equipment, Materials, and Reagents

  - Equipment

    - Hamilton Microlab 625 Dilutor with 250 μl and 2500 μl and Hamilton BFP Syringes located in a Biological Safety Cabinet

  - Materials

    - Headspace sample vials with caps

    - Liquid waste container, e.g., 250 ml beaker

    - Lint free wipes

    - Reservoir for the internal standard solution

    - Deionized water

  - Reagents

    - BAC Internal Standard Solution, refer to the DWI Blood Chemistry Unit Technical Procedure for Determination of Alcohol and Acetone in Blood by Headspace Gas Chromatography

- Performance Verification

  - Upon initial setup perform the Hamilton Microlab 600 Qualification Procedure, refer to11.3.

  - Complete a successful daily QCC as described in the DWI Blood Chemistry Unit Technical Procedure for Determination of Alcohol and Acetone in Blood by Headspace Gas Chromatography.

  - Maintain the performance verification results in the DWI Blood Chemistry Unit.

- Maintenance

  - Record any maintenance in the Hamilton Dilutor Log.

    - Replace the syringes annually at the time of annual calibration.

  - A successful quality control check must be performed prior to analyzing samples, refer to 8.

- Calibration

  - The instrument shall be calibrated annually after the first year of service by an approved vendor.

  - Record calibrations in the Hamilton Dilutor Log and maintain calibration certificates in the DWI Blood Chemistry Unit.

- Quality Control

  - Refer to the DWI Blood Chemistry Unit Technical Procedure for Determination of Alcohol and Acetone in Blood by Headspace Gas Chromatography for quality control check procedures.

- Procedure

  - Cover or seal liquids containing volatile substances to prevent evaporation of the volatiles.  Fill a reservoir with the BAC internal standard solution and place the left diluent tube in the internal standard solution.  Do not seal or cover the reservoir tightly - a vacuum may form and interfere with functioning.

  - Turn power on with the switch on the front of the unit.

  - Selections on the display screen may be entered by pressing the icon on the screen.  Values may be entered via the numeric keypad that appears on the screen.  Arrow keys on the screen may be used to move from one data entry field to another.

  - Method programming.  The BAC Method is initially programmed as follows and may be adjusted to accommodate instrument performance.  Document any adjustments in the Hamilton Dilutor log and perform a successful calibration and daily QCC according to the DWI Blood Chemistry Unit Technical Procedure for Determination of Alcohol and Acetone in Blood by Headspace Gas Chromatography prior to analyzing samples.  Entries to be selected are in bold face type.  It is not necessary to program the method for each use, proceed to 9.5.

    - Configuration – Syringe Set-up

      - Syringe Size			Left – 2500		Right – 250

      - Syringe Flow Rate		Left – 625		Right – 125

      - Initialize Flow Rate		Left – 625		Right – 125

      - Return Steps			Left – 24			Right – 24

      - Back-Off Steps			Left – 96			Right – 80

    - Configuration - Valve Type	Dual Diluter

    - Configuration – Trigger		Either

    - Wizards – Dilution

      - Left Syringe			1800

      - Right Syringe			  200

      - Air Gap			    15

      - Refill				   On

      - Final Volume			2000

    - Wizards – Advanced

      - Air Gap Mode			Auto

      - Air Gap Delay			1.5 S

      - Wash Mode 			Off

      - Program Mode			Diluent and Sample

      - Fill Speed			Left – 625		Right – 100

      - Aspirate Speed			Left – 312		Right – 100

      - Dispense Speed		Left – 625		Right – 100

      - Probe Light Green		Trigger

      - Probe Light Red		Never

  - Using the BAC Method

    - Ensure the left diluent tube is in the internal standard solution.

    - Prime the system by pushing the Prime button on the front of the instrument.

    - Prime the system several times and ensure that the air has been purged from the tubes.

    - Select the Favorites icon, then the Dilution icon and then select BAC.

    - The following steps can be activated by pressing the Run/Pause key or by pressing the trigger button on the hand probe.

    - Collect an aliquot of the sample to be analyzed by placing the tubing on the hand probe into the sample and aspirating the sample by pressing the trigger button.  The sample is contained in the tubing and should never be aspirated into the syringe.  An air gap separates the sample in the tubing from the internal standard solution, preventing the two from mixing.

    - When the submitted sample size allows, discard this aliquot by dispensing to a waste container.

    - Wipe the outside of the tubing with a clean lint free wipe to remove any residue.

    - Collect an aliquot of the sample to be analyzed by placing the tubing on the hand probe into the sample and aspirating the sample be pressing the trigger button.

    - Wipe the outside of the tubing with a clean lint free wipe to remove sample residue.

    - Dispense the sample aliquot and diluent into a headspace vial, labeled to correspond to the sample, by pressing the trigger button. The sample is dispensed first, followed by the internal standard solution, which flushes the sample from the tubing.  After the sample and diluent are dispensed into the head space vial, the left syringe automatically refills with the internal standard solution and an air gap is created in the tubing.

    - Seal the vials by screwing the top onto the headspace vial.

    - Wash the tubing between each sample by aspirating a minimum of one deionized water sample and then dispensing the water/internal standard solution into a waste container.

    - Flush the tubing with a mixture of approximately 10% Clorox and deionized water or equivalent daily, after use, to remove protein build-up and prevent bacterial growth in the tubing.

- Safety

  - Refer to the CCBI Health and Safety Manual.

- References

  - Hamilton Microlab 600 Series Basic Manual, April 2010, Hamilton Company, Nevada, USA.

  - Hamilton Microlab 600 Series Basic Manual, March 2011, Hamilton Company, Nevada, USA.

  - Hamilton Microlab 600 Qualification Procedure, Hamilton Company, Nevada, USA.

- Records

  - Hamilton Dilutor Log

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
|  |  | Updated section 3. |
| 1/16/15 | 2 | Correct reference in section 1 and included reference to maintenance. Include requirement for annual syringe replacement. Include instruction for aspirating and discarding an initial aliquot when sample size allows. Require a minimum of one rinse with deionized water between samples. |
| 4/10/2018 | 3 | Updated safety manual name in 10. |

# 6: Technical Procedure for Biotage TurboVap LV Evaporator

- Purpose / Scope - This procedure provides direction for the use of the Biotage TurboVap LV Evaporator to concentrate extracted samples in the DWI Blood Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification.

- Definitions

  - N/A

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

- Equipment and Materials

  - Equipment

    - Biotage TurboVap LV Evaporator located in a fume hood with a 110/120 VAC power supply

  - Materials

    - Nitrogen, ultra high purity grade

    - Deionized water

    - 16 x 125mm or 13 x 100mm glass test tubes

    - 16 x 125mm or 13 x 100mm TurboVap LV test tube rack

    - Clear bath or equivalent antibacterial agent

- Procedure

  - Record any maintenance or comments in the TurboVap Log.  The addition of water and antibacterial agent need not be recorded.

  - Fill water bath, if necessary, with deionized water and antibacterial agent.

  - Place the appropriately sized test tube rack into the water bath.

  - Power on the TurboVap with the switch at lower left-hand side of instrument.

  - Set water bath temperature with the TEMP numeric push-wheel and allow the water to reach the set temperature.  Water temperature is stable when the TEMP light stops blinking.

  - Place the test tubes containing extracted samples, up to 50, into the rack in the water bath.

  - Set the TIME push-wheel to the desired drying time, e.g., 10 ml of methylene chloride, at  35 degrees C, at 12 psi of nitrogen, will evaporate in approximately 20 minutes.

  - Select the row(s) containing the extracted sample(s) by pressing its corresponding TUBE         STATIONS push-button, choose the button that aligns horizontally with the row.

  - Turn on the nitrogen and set the gas pressure at 10-12 psi with the regulator adjustment valve located on the left side of the instrument.

  - Check the fume hood display and ensure that no error messages are displayed.  If there is an error do not use the evaporator and notify the DWI Blood Chemistry Technical Leader for scheduling of fume hood service.

  - Press the START button.

  - Monitor the samples and remove the test tubes immediately upon evaporation of the solvent.  Increase the time period setting on the instrument if needed.  The TurboVap will sound an alarm when the set time period has expired.

- Safety

  - Refer to the CCBI Health and Safety Manual

- Records

  - TurboVap log

- References

  - TurboVap LV Evaporator Users Manual, 2010, Biotage Corporation, Charlotte, NC.

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
|  |  | Updated section 3. |
| 1/16/15 | 2 | Corrected reference to technical leader in 5.10 |
| 4/10/2018 | 3 | Updated safety manual name in 6. |

# 7: Technical Procedure for Determination of Alcohol, Acetone, and Other Volatiles in Blood by Headspace Gas Chromatography

- Purpose / Scope - This procedure specifies the required elements for calibration and use of the headspace gas chromatograph to determine alcohol (ethanol, methanol, and isopropanol) and acetone concentration in blood in the DWI Blood Chemistry Unit of the CCBI Crime Laboratory.

- Definitions

  - Quality control check – Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

  - Performance verification – The initial confirmation of the reliability of a previously or externally validated method or instrument.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

  - BAC – Blood alcohol concentration

  - GC – Gas chromatograph

  - QCC – Quality control check

- Equipment, Materials and Reagents

  - Equipment and Materials

    - Agilent 7890 Gas chromatograph equipped with flame ionization detectors with Agilent or equivalent DB-ALC1 (front) 30m x 0.535mm  x 3.00µm and DB-ALC2 (back) 30m x 0.535mm  x 2.00µm capillary columns, G6509 headspace autosampler, data station and printer.

    - Hamilton-Microlab 625 dilutor

    - Mettler XS1003S Balance

  - Materials

    - Headspace vials with sealing caps

    - Deionized water

    - Helium gas, UHP Grade

    - Hydrogen gas, UHP Grade

    - Nitrogen gas, UHP Grade

    - Compressed Air, Zero Grade

  - Reference Materials

    - n-Propanol (1-propanol), Certified Grade or higher

    - Multi-component solutions, NIST traceable, containing ethanol, methanol, acetone and isopropanol.  Concentrations - 0.010 g / 100 mL, 0.050 g/100 mL, 0.100 g/100 mL and 0.400 g/100 mL.

    - Multi-component solutions, NIST traceable, containing ethanol, methanol, acetone and isopropanol.  Concentrations - 0.010 g / 100mL, 0.025 g / 100 mL, 0.050 g/ 100 mL, 0.100 g/ 100 mL, 0.200 g / 100 mL, 0.400 g/ 100 mL.

  - Prepared Reagents

    - BAC Internal Standard Solution (n-propanol)

      - Prepare a 0.050 gram / 100 mL solution of n-propanol reference material.

Example: weigh 1.00 gram of n-propanol into a beaker.  Quantitatively transfer to a 2000 mL volumetric flask.  Bring the flask to volume with deionized water.

      - Lot Number:  Eight digit format year/month/day/BacIS/initials of preparer.

Example:  20130101BacISXXX

      - Expiration:  Six months.

      - Storage:  Refrigerate.

      - QCC:  Successful calibration, refer 8.

- Gas Chromatograph Performance Verification for New Instruments

  - New gas chromatographs shall be installed by a manufacturer representative and shown to meet manufacturer requirements.

  - Complete a performance verification on new gas chromatographs prior to use for casework.  Maintain the data in the DWI Blood Chemistry Unit.

  - Include in the performance verification:

    - Successful calibration, refer to 8.

    - Successful daily QCC, refer to 9.1.1.

    - Analysis of six preparations of multi-component reference material from 4.3.3. at concentrations of 0.010, 0.025, 0.050, 0.100, 0.200 and 0.400 g / 100 mL, over at least a two day period.  All quality control requirements must be met.

      - The % RSD at each concentration must be 2% (5% for acetone) or less.

      - The individual concentration results must be within ± 5% (10% for acetone) of the target concentration for concentrations of 0.040 g / 100 mL or greater.  For samples with target concentrations of less than 0.040 g / 100 mL, the individual concentration results must not differ from the target concentrations by more than ± 0.002 g / 100 mL.

    - The analysis of three preparations of reference material at 0.004 g / 100 mL

      - All components must be identified.

    - The analysis of five preparations of 0.080 g / 100 mL multi-component samples prepared using negative blood (containing sodium fluoride preservative and oxalate anticoagulant) and five preparations of negative blood.

      - The % RSD at each concentration must be 2% (5% for acetone) or less.

      - The individual concentration results must be within ± 5% (10% for acetone) of the target concentration.

- Gas Chromatograph Maintenance

  - Record all maintenance in the maintenance log at the time it is performed.

  - When maintenance is performed, the instrument shall be out of service until the required daily system check and/or calibration is successfully completed and recorded in the instrument log.  File any generated data in the instrument logbook maintained near the instrument.

  - Suggested Routine Maintenance Schedule

    - This is a suggested maintenance schedule.  Instrument use may alter the need for maintenance.  The maintenance schedule shall be determined by the DWI Blood Chemistry Unit Technical Leader based upon instrument usage.

    - Septum

      - Replace weekly when in use.

      - A successful daily QCC, refer to 9.1.1, shall be performed prior to analyzing samples.

    - Syringe

      - Inspect every six months, replace as needed.

      - A successful daily QCC, refer to 9.1.1, shall be performed prior to analyzing samples.

    - Liner

      - Inspect every twelve months, replace as needed.

      - A successful daily QCC, refer to 9.1.1, shall be performed prior to analyzing samples.

    - Jet

      - Inspect every twelve months, replace as needed.

      - A successful daily QCC, refer to 9.1.1, shall be performed prior to analyzing samples.

    - Column

      - Replace if chromatographic deterioration is noted.

      - A successful calibration, refer to 8, and daily QCC, refer to 9.1.1, shall be performed prior to analyzing samples.

    - Non-routine Maintenance

      - All non-routine maintenance shall, at a minimum, be followed by a successful daily QCC, refer to 9.1.1, prior to analyzing samples.

      - Non-routine maintenance shall be evaluated by the DWI Blood Chemistry Unit Technical Leader to determine the need for recalibration prior to analyzing samples.

    - Shutdown

      - Successful daily QCC, refer to 9.1.1, shall be performed following any GC or autosampler shutdown.

      - Record the shutdown in the maintenance log.

- Gas Chromatograph Parameters

  - The stated method parameters are initial settings that may be adjusted to accommodate instrument performance.  Document any adjustments in the instrument log and perform a successful calibration, refer to 8, and daily QCC, refer to 9.1.1, prior to analyzing samples.

  - Gas Chromatograph Method Parameters

    - Isothermal Column Temperature:  35 °C

    - Injector Temperature:  200 °C

    - Detector Temperature:  200 °C

    - Column Flow:  9 mL/min, nominal, constant flow

    - Split ratio:  2:1

    - Run Time:  5:00 minutes

    - Detector make up flow: 18 mL/min

    - Detector H2 flow:  30 mL/min

    - Detector Air Flow:  300 mL/min

  - Headspace Autosampler Method Parameters

    - Incubation Temperature:  70 °C

    - Incubation Time:  480 seconds

    - Agitator Speed:  250 rpm

    - Agitator On/Off:  5 seconds on / 30 seconds off

    - Runtime 5:00 minutes

    - Syringe temperature 75 °C

    - Sample volume 1000 µL

    - Syringe fill speed:  100 μL / second

    - Pull-up delay:  2 seconds

    - Injection speed:  700 μL / second

    - Pre-inject delay / Post-inject delay:  0.5 second

    - Flush time:  120 seconds

  - Data Station Method Parameters

    - The peak retention times may be updated daily, when in use, to the retention times of the last injection of the 0.100 g/100 mL NIST traceable multi-component solution in the current Daily QCC.

    - The peak retention time windows shall be set to the peak retention time ± 0.050 minutes + 0.5 % of the peak retention time.

- BAC Calibration

  - Calibrate the GC upon preparation of a new lot of Internal Standard solution and after instrument maintenance that may affect the calibration.

  - Prepare each of the NIST traceable multi-component solutions (0.010, 0.025, 0.050, 0.100, 0.200, 0.400 g / 100 mL) in duplicate according to 9.2.

  - Chromatograph the calibration samples on the gas chromatograph

  - Update the calibration with the calibration samples.

    - The response at each concentration shall be determined by the average response of the duplicates analyzed at that concentration.

    - The retention time for each component shall be that of the final calibration sample injected.

    - The calibration curve shall be fitted to a linear model with equal weighting and with the origin included.

    - The calibration curves for each component shall show a coefficient of determination, r2, of 0.995 or greater.  If the calibration has a coefficient of determination of less than 0.995, take appropriate action, e.g., maintenance and / or preparation of new solution(s), and repeat the calibration.  Record all calibrations and maintenance in the instrument logbook.

  - Save the data analysis method according to the format “BacCal” and eight digit format year/month/day.

Example:  BacCal20130101

  - Reprocess the calibration samples with the updated method.

    - Each component must be identified by the instrument software on both columns and be visually baseline resolved.

    - For those with target concentrations of 0.040 g / 100 mL or greater, each component’s individual quantitation results must be within +/- 5.0 % (10% for Acetone and Methanol) of the target.

    - For those with target concentrations of less than 0.040 g / 100 mL, each component’s individual quantitation results must be within +/- 0.002 g/100 mL (0.004 g/100 mL for Acetone and Methanol) of the target concentration.

  - If the calibration is unacceptable for ethanol, take appropriate action, e.g., maintenance or new solution preparation, and repeat the calibration.  If the calibration is unacceptable for a component other than ethanol either take appropriate action, e.g., maintenance or new solution preparation, and repeat the calibration or record the unacceptable component on the calibration data and in the calibration log.  Do not perform quantitations of the unacceptable component.

  - BAC Calibration Verification

    - Prepare each of the NIST traceable multi-component solutions (0.010, 0.050, 0.100 and 0.400 g / 100 mL) in duplicate according to 9.2.

    - Chromatograph the verification samples on the gas chromatograph.

    - Quantitate the verification samples with the updated calibration.

      - Each component must be identified by the instrument software on both columns and be visually baseline resolved.

      - For the duplicates with target concentrations greater than 0.04 g/100 mL, evaluate each component’s quantitation results according to the equation

Range tolerance =   100 × [(highest measured concentration) – (lowest measured concentration)]

(lowest measured concentration)

The range tolerance must be ≤ 5.0 % (10 % for methanol and acetone).  Any components that do not meet this requirement will not be quantitated.

        - Each component’s individual quantitation results must be within ± 5.0 % (10 % for methanol and acetone) of the target concentration.  Any components that do not meet this requirement shall not be quantitated.

      - For the 0.01 g / 100 mL duplicates, each component’s individual quantitation results must be within +/- 0.002 g / 100 mL (0.004 g/ 100mL for Acetone and Methanol) of the target concentration.  Any components that do not meet this requirement shall not be quantitated.

    - If the calibration is unacceptable for ethanol, take appropriate action, e.g., maintenance or new solution preparation, and repeat the calibration.  If the calibration is unacceptable for a component other than ethanol either take appropriate action, e.g., maintenance or new solution preparation, and repeat the calibration or record the unacceptable component on the calibration data and in the calibration log.  Do not perform quantitations of the unacceptable component.

  - Prepare a file labeled with the name of the calibrated method, refer to 8.5.  Maintain this file on the shared drive in the DWI Blood Chemistry Unit folder.  This Calibration file shall include:

    - Reprocessed calibration sample instrument printouts.

    - Verification sample instrument printouts

    - Lot numbers of the prepared reagent:  internal standard solution

    - Lot numbers of the NIST traceable multi-component solutions

    - Calibration table and the r2 values of each component.

  - Record each calibration in the calibration log with the date, lot number of internal standard used, internal standard expiration date and operator initials.

- Procedure

  - Quality Control Checks

    - Daily Quality Control Check

      - The daily quality control check must be performed in the twenty-four hours preceding any sample analysis.  The daily quality control check is valid for twenty four hours after the injection of the last verification standard.

      - Prepare each of the NIST traceable multi-component solutions, 0.010, 0.050, 0.100 and 0.400 g/100mL, in duplicate according to 9.2.  Prepare a water blank in duplicate using deionized water according to 9.2.

      - Chromatograph the prepared samples in a single sequence with a single water blank at the beginning followed by the NIST traceable duplicates from low to high concentration and a single water blank after the 0.400 g/100 mL NIST traceable multi-component solution samples.

      - Quantitate the samples with the current data analysis method corresponding to the internal standard used.

        - Each NIST traceable multi-component solution sample must meet the requirements stated in 8.8.3.1 through 8.8.3.3.

        - The water blank must not contain any identifiable methanol, ethanol, isopropanol, acetone or any other identifiable volatile.

      - If the QCC is unacceptable for ethanol, take appropriate action, e.g., maintenance or new solution preparation, and repeat the QCC.  If the QCC is unacceptable for a component other than ethanol either take appropriate action, e.g., maintenance or new solution preparation, and repeat the QCC or record the unacceptable component on the QCC data and in the instrument log.  Do not perform quantitations of the unacceptable component

      - Prepare a file labeled “BACQCC” followed by the date in yyyymmdd format.  Maintain this file on the shared drive in the DWI Blood Chemistry Unit folder.  This QCC file shall include:

        - NIST traceable multi-component solution sample instrument printouts

        - NIST traceable multi-component solution lot numbers

        - Water blank sample instrument printouts

        - Lot number and expiration of internal standard solution

      - Record each daily QCC in the instrument log with initials, date, time of the first injection, lot number of internal standard used, internal standard expiration date and operator initials.

    - In-Sequence Quality Control Checks

      - A sample sequence must contain a minimum of 10 % control samples.

      - The sequence must contain at least one positive control sample and one negative control sample, each prepared in duplicate.

      - The first and last samples of a sequence shall be control samples with any remaining required control samples evenly distributed throughout the batch.  The samples and controls should be prepared in the same order in which they will be analyzed.

      - Select the target concentration of the positive control samples to avoid duplicate concentrations within a sequence when possible.

      - Negative control

        - Prepare a deionized water sample according to 9.2.

        - Negative controls must not contain any identifiable methanol, ethanol, isopropanol or acetone.

      - Positive control

        - Prepare the number and concentrations needed of the NIST traceable multi-component solutions, refer to 4.3.6.

        - Each positive control sample component must meet the requirements in 8.8.3.1 through 8.8.3.3.

      - Record any in-sequence QCC sample components that do not meet all requirements in the instrument log and follow 9.1.2.7.1 or 9.1.2.7.2.

        - If an in-sequence QCC sample is found to be unacceptable for ethanol, all cases in the sequence after the last acceptable positive quality control sample shall be reanalyzed.  This shall be documented in the case file.  Data generated after the last acceptable positive control sample shall not be used to identify or quantify ethanol in casework.

        - If an in-sequence QCC sample is found to be unacceptable for a component other than ethanol, cases in the sequence after the last acceptable positive quality control sample for that component shall be reanalyzed if that component is present.  This shall be documented in the case file.  Data generated after the last acceptable positive control sample shall not be used to identify or quantify that component in casework.

        - Record each sequence in the instrument log with initials, date and time of the first injection, internal standard solution lot number and expiration date.

      - Include the in-sequence QCC data in the “BACQCC” file from 9.1.1.6.  The file shall include:

        - Sequence table

        - Positive and negative control instrument printouts

        - Lot numbers of the BAC working verification solutions and / or NIST traceable multi-component solutions

  - Alcohol and Acetone Concentration in blood

    - Sampling

      - Allow all solutions and samples to be analyzed to equilibrate to room temperature.

      - Ensure that blood samples are homogenous by shaking and/or vortexing.

        - If a homogenous sample cannot be obtained due to the presence of clots, a notation shall be made in the case file and the alcohol concentration shall be calculated according to 9.2.7.6.

        - If a homogenous sample cannot be obtained because the blood cells have been separated from the liquid, a notation shall be made in the case file and the alcohol concentration shall be calculated according to 9.2.7.6.

        - If a homogenous sample cannot be obtained for any other reason, a notation shall be made in the case file detailing the condition of the sample and its handling.

    - With the Hamilton Microlab 625 Dilutor, refer to the DWI Blood Chemistry Unit Technical Procedure for Use of the Hamilton Microlab 625 Dilutor to Prepare Samples for Blood Alcohol Determination, deliver 1.80 mL of the BAC Internal Standard Solution and 0.20 mL of the liquid to be analyzed into a headspace vial and cap securely.

    - Prepare each liquid to be analyzed in duplicate and chromatograph using the BAC method.

    - Quantitate each sample using the current calibrated BAC method corresponding to the lot of BAC Internal Standard Solution used in sample preparation.

    - The value of the multiplier in the sequence table shall be 10.

    - Identification of alcohol, acetone, and other volatiles

      - Ethanol, methanol, isopropanol or acetone shall be integrated in the appropriate retention time window on both columns in both sample preparations to be identified.

      - Include the sample instrument printouts, calibration file reference and quality control check file reference in the case file.

      - Volatiles other than ethanol, methanol, isopropanol or acetone may be qualitatively identified using a reference material relative retention time (RRT) comparison relative to the internal standard.

        - The reference material shall be analyzed in the same sequence and under the same chromatographic conditions as the sample.

        - The RRT of the sample and the reference material shall not differ by more than 1.0% on each column in both sample preparations.

        - Include the sample chromatograms, the sample RRT, the reference material RRT, calibration file reference and quality control check file reference in the case file.

      - Reporting

        - Refer to the DWI Blood Chemistry Unit Technical Procedure for DWI Blood Chemistry Analysis for reporting of identified alcohols and acetone.

    - Determination of alcohol or acetone concentration

      - The mean of the four measured values obtained for each component is the concentration of that component.

      - The Blood Alcohol Concentration is the sum of the means of the concentrations of the identified alcohols in a blood sample.  Each mean shall be truncated to the hundredths place prior to summation.

      - The concentration of acetone shall be truncated to the hundredths place.

      - Reanalyze samples with identified alcohol or acetone values greater than or equal to 0.040 g / 100 mL that exhibit greater than 5 % (10 % for methanol and acetone) range tolerance as described by the equation:

Range Tolerance =   100 × [(highest measured concentration) – (lowest measured concentration)]

(lowest measured concentration)

      - Reanalyze samples with identified alcohol or acetone values less than 0.040 g / 100 mL that exhibit a difference between the highest and lowest measured concentrations greater than 0.004 g / 100 mL.

      - Reanalysis shall be performed a minimum of 1 time.  The nature of the sample, sample volume, and specific details of the case should be considered.

        - If the sample meets the criteria above when reanalyzed, report using the 4 concentrations obtained during reanalysis.

        - If the sample does not meet the criteria above when reanalyzed, report the average of all the concentration measurements, the lowest measurement, the highest measurements, along with a statement address why the uncertainty is not determined. (i.e. Due to the nature of the sample, uncertainty of measurement was not determined)

      - Clotted blood samples that cannot be rendered homogenous and samples in which the blood cells have been separated from the liquid (including serum and plasma) shall be converted to an equivalent whole blood alcohol concentration by dividing the alcohol concentration by 1.2  to compensate for the whole blood:serum alcohol distribution ratio.

      - Reporting

        - Refer to the DWI Blood Chemistry Unit Technical Procedure for DWI Blood Chemistry Analysis for reporting of identified alcohols and acetone.

- Calculations

      - Alcohol or acetone values range tolerance:

Range tolerance =   100 × [(highest measured concentration) – (lowest measured concentration)]

(lowest measured concentration)

      - Conversion of non-homogenous samples and samples in which the blood cells have been separated from the liquid to an equivalent whole blood alcohol concentration:  divide the alcohol concentration by 1.2.

- Uncertainty of Measurement

  - Refer to the DWI Blood Chemistry Unit Technical Procedure for Uncertainty of Measurement.

- Limitations

  - Clotted samples that cannot be rendered homogenous and samples in which the blood cells have been separated from the liquid (including serum and plasma) shall be converted to an equivalent whole blood alcohol concentration by dividing the measured alcohol concentration by 1.2 to compensate for the whole blood:serum (plasma) alcohol distribution ratio.

  - Refer to the Determination of Alcohol and Acetone in Blood by Headspace Gas Chromatography validation data in the DWI Blood Chemistry Unit for limit of detection and limit of quantitation.

  - Quantitation values greater than the highest calibration standard, 0.40 g / 100 ml, may be diluted and reanalyzed or reported as “greater than 0.40 gram per 100 ml.”

  - Volatiles other than alcohol and acetone are only determined qualitatively.

- Safety

  - Refer to the Health and Safety Manual.

- References

  - Randall C. Baselt.  Disposition of Toxic Drugs and Chemicals in Man. 8th Ed. (2008): 561 – 565.

  - Agilent 7890A GC User Information, Agilent Instrument Utilities Version B.1.06.11343.1852.

  - Agilent GC Sampler 80/120 User Manual, Edition 03/2010, Agilent Technologies.

  - James C. Garriott (Editor), Medicolegal Aspects of Alcohol, 5th  Ed., 2008.

  - Moffat, Anthony C. Clarke's Analysis of Drugs and Poisons: In Pharmaceuticals, Body Fluids and Postmortem Material, 4th Ed., London: Pharmaceutical Press, 2004. Print.

  - G. Machata, "Determination of Alcohol in Blood by Gas Chromatographic Head Space Analysis," Clinical Chemistry Newsletter, Vol. 4 No. 2 (1972), 29-32.

  - Static Headspace Blood Alcohol Analysis with the G1888 Network Headspace Sampler, Agilent Technologies, Inc., 2004, Publication Number 5989-0959EN.

  - T. Macchia, et al, "Ethanol in Biological fluids:  Headspace GC Measurement," Journal of Analytical Toxicology, Vol. 19 (1975), 241-245.

  - G. Machata, “The Advantages of Automated Blood Alcohol Determination by Head Space Analysis,” International Journal of Legal Medicine, Vol. 75 (1975), 229-234.

  - Agilent OpenLAB CDS ChemStation Edition Concepts and Workflows, Agilent Technologies Inc., 2011, Manual Part Number M8301-90012.

  - Caplan, Yale H. and Goldberger, Bruce A. (Editors), Garriott’s Medicolegal Aspects of Alcohol, 6th Ed., 2015.

- Records

  - Case file

  - Quality control check file

  - Calibration file

  - GC Instrument log

  - GC Maintenance log

  - GC Calibration log

  - Reagent log

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB documents |
|  |  | Updated section 3. Added wording to line 4.4.4.1. and reference 14.10. |
| 1/16/15 | 2 | Updated for minor additions, deletions, corrections and clarifications: 4.1.5, 4.4.3, 4.4.4, 7.4.1, 8.8.3.2, 9.1.1.2, 9.1.1.3, 9.1.1.4.1, 9.1.2.2, 9.1.2.3, 9.1.2.6.2, 10.1.1.2, 12.1, 14.11 |
| 11/3/16 | 3 | Made methanol criteria for QCC compliance identical to criteria for acetone in: 8.8.3.2, 8.8.3.2.1, 9.2.7.4 |
| 2/14/17 | 4 | Updated maintenance requirements in 6.3.3.1, 6.3.4.1, and 6.3.6.1 |
| 3/2/18 | 5 | Updated BAC Internal Standard Expiration Date and Storage in 4.4.3.3 and 4.4.3.4, updated 8.6.2 to be consistent with verification requirement, updated 9.1.2.7. |
| 4/10/2018 | 6 | Changed time frame for Daily QCC in 9.1.1.1. Corrected spelling 9.2.6.2. Updated safety manual name in 13. |
| 10/07/19 | 7 | Removed volumetric flasks and pipettes from 4.1 Equipment and Materials. Removed Ethanol, Isopropanol, Acetone, and Methanol from 4.3 Reference Material. Added Multi-component solutions to 4.3 Reference Material. Removed BAC Stock and Working Calibration Solutions from 4.4 Prepared Reagents. Revised 5.3.3. Added NIST traceable multi-component solutions to 8.2 BAC Calibration. Removed calibration solutions from 8.9.3. Revised 5.12, 8.9, and 9.1.1.6. to allow for electronic storage. Added 9.2.7.6. to address reanalysis. Changed value of highest calibration standard in 12.3. |
| 10/25/19 | 8 | Changed criteria for acetone and methanol in 8.6.3 and 8.8.3.3. Added 9.7.6.1 and 9.7.6.2 to address reanalysis. |
| 5/14/20 | 9 | Title: Added other volatiles 8.8. Calibration Verification - changed % Diff to range tolerance 9.2.6. Identification of alcohol and acetone - added other volatiles. Added 9.2.6.3. for other volatile qualitative analysis criteria and requirements 9.2.7.4. Determination of alcohol or acetone concentration - changed % difference to range tolerance 10. Calculations - changed % difference to range tolerance 12. Limitations - added limitative 12.4 regarding qualitative identification of other volatiles |

# 8: Technical Procedure for Enzyme Linked Immunosorbent Assay (ELISA) as a Drug Screen

- Purpose / Scope – This procedure details the use of the Tecan Workstation and Enzyme Linked Immunosorbent Assay to screen blood for controlled substances, controlled substance metabolites and non-controlled substances in the DWI Blood Chemistry Unit of the CCBI Crime Laboratory.

- Definitions

  - Quality control check – Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

  - Performance verification – The initial confirmation of the reliability of a previously or externally validated method or instrument.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

- Equipment, Materials and Reagents

  - Equipment

    - Tecan EVO 75/2 liquid handling platform

    - Tecan HydroFlex washer

    - Tecan Sunrise reader

    - Data station

  - Materials

    - Vortex mixer

    - Test tubes, 12 x 75mm

    - Test tube stoppers or caps

    - Class A volumetric flask, 10 ml

    - Mechanical pipetters: 0.010 ml, 0.100 ml, 0.20 ml, 0.250 ml, and 1.00 ml

    - Deionized water

  - Reference Materials

    - (±) Amphetamine, 1 mg/ml - Diluted 1:10 with methanol for a concentration of 0.1 mg/ml

    - Benzoylecgonine, 1 mg/ml

    - Delta-9-carboxy-11-nor-delta-9-tetrahydrocannabinol (THC-COOH), 1 mg/ml - Diluted 1:10 with methanol for a concentration of 0.1 mg/ml

    - Fentanyl, 1 mg/ml, - Diluted 1:100 with methanol for a concentration of 0.01 mg/ml

    - (±)-Methadone, 1 mg/ml

    - (±)-Methamphetamine, 1mg/ml - Diluted 1:10 with methanol for a concentration of 0.1 mg/ml

    - Morphine, 1 mg/ml – Diluted 1:10 with methanol for a concentration of 0.1 mg/ml

    - Nordiazepam, 1 mg/m – Diluted 1:10 with methanol for a concentration of 0.1 mg/ml

    - Phenobarbital, 1 mg/ml

  - Critical Reagents

    - Negative blood

  - Commercial Reagents

    - Methanol, ACS grade

    - 1 N HCl, ACS grade

    - 1 N NaOH, ACS grade

    - Immunalysis ELISA kits may include amphetamine, barbiturates, benzodiazepines, cocaine  metabolite, metabolites of delta-9-THC, fentanyl, methamphetamine, methadone, and opiates containing:

      - 96 Well - coated microplates

      - Enzyme conjugate, matched to the microplates

      - TMB substrate reagent

      - Stop reagent

  - Prepared Reagents

    - ELISA Calibration Solution

      - To a 10 ml class A volumetric flask, add the applicable Reference Materials:

        - 0.025 ml Benzoylecgonine, 1 mg/ml

        - 0.050 ml Delta-9-carboxy-11-nor-delta-9-tetrahydrocannabinol, 0.1

mg/ml

        - 0.025 ml (±)-Methadone, 1 mg/ml

        - 0.100 ml (±)-Methamphetamine, 0.1 mg/ml

        - 0.050 ml Morphine, 0.1 mg/ml

        - 0.050 ml Nordiazepam, 0.1 mg/ml

        - 0.150 ml Phenobarbital, 1 mg/ml

        - 0.100 ml (±) Amphetamine, 0.1 mg/ml

        - 0.050 ml Fentanyl, 0.01 mg/ml

      - Dilute the flask to volume with methanol.

      - Lot number:  yyymmdd / ELISACal / Initials

      - Expiration:  See DWI Blood Chemistry Technical Procedure for Quality Assurance.

      - Storage:  freezer

      - QCC:  refer to Quality Control Check.

    - ELISA Verification Solution

      - Prepare with Reference Materials from a different supplier or lot than those used

in 4.6.1

      - To a 10 ml class A volumetric flask, add the applicable Reference Materials:

        - 0.025 ml Benzoylecgonine, 1 mg/ml

        - 0.050 ml Delta-9-carboxy-11-nor-delta-9-tetrahydrocannabinol, 0.1

mg/ml

        - 0.025 ml (±)-Methadone, 1 mg/ml

        - 0.100 ml (±)-Methamphetamine, 1 mg/ml

        - 0.050 ml Morphine, 0.1 mg/ml

        - 0.050 ml Nordiazepam, 1 mg/ml

        - 0.150 ml Phenobarbital, 1 mg/ml

        - 0.100 ml (±)-Amphetamine, 0.1 mg/ml

        - 0.050 ml Fentanyl, 0.01 mg/ml

      - Dilute the flask to volume with methanol.

      - Lot number:  yyyymmdd / ELISAVer / Initials

      - Expiration:  See DWI Blood Chemistry Technical Procedure for Quality Assurance.

      - Storage:  freezer

      - QCC:  refer to Quality Control Check.

- Performance Verification for New Instruments

  - New Tecan workstations shall be installed by a manufacturer representative and shown to meet any manufacturer’s requirements.

  - Performa a performance verification on new Tecan workstations prior to use for casework according to the CCBI Crime Laboratory Administrative Procedure for Validations and Performance Verifications.

  - Maintain the performance verification documentation in the DWI Blood Chemistry Unit.

- Maintenance

  - Record all maintenance in the ELISA maintenance log or ELISA monthly maintenance log at the time it is performed.

  - Record any instrument errors or malfunctions in the instrument log at the time they are observed along with any action taken.  Notify the DWI Blood Chemistry Unit Technical Leader of any malfunctions that cannot be corrected and mark the instrument log “Out of Service.”  The Technical Leader shall correct the problem or schedule repair by an approved vendor and mark the instrument log “In Service” when the instrument is returned to service.

  - Preventive maintenance shall be performed by an approved vendor annually.

  - Monthly maintenance

    - Empty and fill the two deionized water reservoirs, flush and prime the system.

    - Perform an acid/base wash of the instrument using 1 N HCl and 1N NaOH.

      - Turn on the computer, workstation, washer, and reader.

      - Open Navitrak 2, Utilities, and select Acid/Base Wash.  Follow the instrument instruction to place the system intake tube into a beaker containing 250 ml of 1N HCl.  The instrument will rinse and soak with the acid solution for 15 minutes.

      - After the 15 minute time period, follow the instrument instruction to place the intake tube into a beaker containing 500 ml of deionized water.  The instrument will rinse with deionized water.

      - Follow the instrument instruction to place the system intake tube into a beaker containing 250 ml of 1N NaOH.  The instrument will rinse and soak with the base solution for 15 minutes.

      - After the 15 minute time period, follow the instrument instruction to place the intake tube into a beaker containing 500 ml of deionized water.  The instrument will rinse with deionized water.

      - Follow the instrument instruction to place the system intake tube into a beaker containing 250 ml of 1N HCl.  The instrument will rinse with the acid solution.

      - Follow the instrument instruction to place the system intake tube into the system deionized water container and to conduct a flush / prime of the system.

      - When the flush / prime has completed, exit Navitrak 2 and direct the instrument to return the robotic arms to the home position.

      - Turn off the workstation, reader and computer.  Select Night Rinse on the washer.

- Procedure

  - Allow all solutions and samples to equilibrate to room temperature.

  - Check system liquid containers and fill with deionized water if necessary.  Check the waste containers and empty if necessary.

  - Turn on the Tecan workstation, washer, reader and computer in that order.

  - Open Navitrak 2, Utilities, and select Flush/Prime. The instrument will position the robotic arms and prime / flush the two pipets.  When completed, exit Navitrak 2 and direct the instrument to return the robotic arms to the home position.  The instrument is now ready to begin analysis of samples.

  - Ensure that all blood samples to be analyzed are homogenous by shaking and/or vortexing.

    - If a homogenous sample cannot be obtained, make a notation in the case file detailing the condition of the sample and its handling.

  - Calibration Standard

    - At the time of case sample preparation, pipet 0.100 ml of the ELISA Calibration Solution into a test tube.

    - Add 4.9 ml of negative blood to the test tube.

    - Cap and vortex the test tube.

    - The concentration of the calibration standard is 300 ng/ml phenobarbital, 10 ng/ml nordiazepam, 10 ng/ml morphine, 50 ng/ml benzoylecgonine, 50 ng/ml methadone, 10 ng/ml THC-COOH, 1 ng/ml fentanyl, 20 ng/ml methamphetamine, and 20 ng/ml amphetamine.

    - Prepare as directed in 7.10.  Dispose of any unused portion.

  - Verification Standard

    - At the time of case sample preparation pipet 0.200 ml of the ELISA Verification Solution to a test tube.

    - Add 4.8 ml of negative blood to the test tube.

    - Cap and vortex the test tube.

    - Prepare as directed in 7.10.  Dispose of any unused portion.

    - The concentration of the verification standard is 600 ng/ml phenobarbital, 20 ng/ml nordiazepam, 20 ng/ml morphine, 100 ng/ml benzoylecgonine, 100 ng/ml methadone, 20 ng/ml THC-COOH, 2 ng/ml fentanyl, 40 ng/ml methamphetamine, and 40 ng/ml amphetamine.

  - Negative Standard

    - Prepare a negative standard as described in 7.10 using negative blood.

  - Pipet 0.250 ml of the calibration standard, verification standard, negative standard and each blood sample to be analyzed into each of two labeled disposable glass test tubes using a mechanical pipet labeled “Blood.”

    - Pipette 2.5 ml of deionized water into each of the test tubes.

    - Cap / stopper the tubes and vortex for approximately 10 seconds.

    - Examine the tubes and remove any excessive bubbles from the surface of the liquid.

  - Arrange the test tubes into the sample racks.  The calibration standard, verification standard, negative standard and case samples must be prepared in duplicate and run concurrently.  The samples must be arranged in the sample racks in correct order:  negative standard duplicates, then calibration standard duplicates, then the verification standard duplicates, followed by the case sample duplicates.

  - Arrange the plates with the desired assays wells.  Plates two through six may contain one or two assays on each plate, depending on the number of samples that are being analyzed.

  - When the plates are arranged, place the plates on the workstation deck.

  - Place the sample racks containing the negative controls, calibration standards, verification standards and case samples on the workstation deck.

  - Fill the appropriate reagent troughs with TMB substrate and stop reagent.

  - Fill the screw-top vials with conjugate and place these vials in the conjugate rack in the appropriate order based upon the arrangement of assays on the plates.

  - Open Navitrak 2, Test Panel, and configure the plates and assays based upon the arrangement of the plates.  Ensure that the placement of the conjugate vials matches the arrangement of the plates.

  - Select Sample List and prepare the sample list, listing in order the negatives, calibration standards, verification standards and case samples.  Ensure that the placement matches the configuration of the sample list.

  - Select Review and ensure that the configuration of the plates, samples and reagents is correct.

  - Select Run EVO to begin the analysis.

    - The workstation automatically pipets the negatives, standards and case samples into the wells on each plate followed by the conjugates.

    - There is a one hour incubation period.

    - The plates are then rinsed with deionized water in the washer.

    - TMB substrate is pipetted into the remaining assay wells.

    - There is a thirty minute incubation period.

    - The stop acid reagent is pipetted into the wells.

    - The plates are moved to the reader and the absorbance measured and the data printed.

  - After the results of each plate and assay are printed, exit Navitrak 2 and direct the instrument to return the robotic arms to the home position.

  - Print the sample list and plate layout.  Select Night Rinse on the washer.  Turn off the workstation and reader.

  - An inverse relationship exists between absorbance and concentration.

  - The calibration standard, the verification standard, the negative standard and the case sample(s) will be run in duplicate. The two runs will be averaged, and the average absorbance of each of the calibration standard, the verification standard, and case sample (B) is divided by the average absorbance of the negative control (B0) and multiplied by 100 to obtain %B/B0.

    - The negative standard value, (i.e., negative cut-off) is the %B/B0 of the negative standard.

    - The positive standard value (i.e., positive cut-off) is the %B/B0 of the calibration standard.

    - The %B/B0 of the verification standard(s) and case sample(s) are evaluated with respect to the positive standard and the negative standard.

  - Quality Control Check

    - The variation coefficient of all duplicate results must be 20% or less.

    - Case samples with variation coefficients greater than 20% must be rerun unless each individual %B/B0 results is evaluated the same with respect to the positive standard and the negative standard.

    - Negative, calibration or verification standards with variation coefficients greater than 20% must be rerun unless each %B/B0 value for the standard is evaluated such that

      - both negative standards must be greater in value than the calibrator standards,

      - both calibrator standards must have a %B/B0 of less than 80%, and

      - both verification standards must be evaluated as positive with respect to the positive standard.

    - If the requirements listed in 7.24.1 through 7.24.3.1 are not met the analysis must be repeated.

  - Record all analyses in the ELISA instrument log with initials, date, ELISAQC file name, refer to 7.28.  Record any assays that do not meet the requirements in 7.24 along with the requirement and the observed value.

  - The data is exported into an Excel file.

    - Save the ELISA template file with the file name from 7.28.

    - Copy the data for each assay from the generated Excel file and paste it into the corresponding assay of the ELISAQC Excel file.

    - Save the file.

    - Print the summary and all sample worksheets.

  - Evaluate case samples that have an average %B/B0 value less than or equal to the positive cut-off as positive.  Evaluate case samples that have an average %B/B0 value equal to or greater than the positive cut-off as negative.

Sample comparison											Evaluated Result

Sample average >%B/B0= Positive cut-off					Negative

Sample average ≤ Positive cut-off							Positive

  - Prepare a file labeled “ELISAQC” followed by the eight digit format date of the analysis “yyyymmdd.”  Multiple analyses on a single date may be indicated be adding unique letters to the end of the file name in alphabetical order.

    - Maintain this file in the DWI Blood Chemistry Unit folder.

    - Include in the file the summary worksheet from 7.26.4 and:

      - Lot numbers of the microplates

      - Lot numbers of the enzyme conjugates

      - Lot number of the TMB substrate

      - Lot number of the stop reagent

      - Lot number of the ELISA calibration solution

      - Lot number of the ELISA verification solution

      - Lot number of the negative blood

      - All sample and standard individual absorbances

      - All sample and standard average absorbances

      - All sample and standard duplicate absorbance standard deviations

      - All sample and standard duplicate absorbance variation coefficients

      - Adjusted negative standard value

    - Reference the file in each sample case file.

  - Record in each sample case file the printed sample worksheet from 7.26.4 and:

    - Reference to the ELISAQC file

    - Lot numbers of the microplates

    - Lot numbers of the enzyme conjugates

    - Lot number of the TMB substrate

    - Lot number of the stop reagent

    - Lot number of the ELISA calibration solution

    - Lot number of the ELISA verification solution

    - Lot number of the negative blood

    - Case sample, %B/B0 average, variation coefficient

    - Case sample evaluation, if evaluated

- Reporting

  - Refer to the DWI Blood Chemistry Unit Technical Procedure for DWI Analysis.

- Calculations

  - Mean, x =xn   where n = number of measured values

  - Standard deviation, s=x-x2n-1

  - Variation coefficient,  CV (%)= 100sx

  - %B/B0, absorbance of the calibration standard, verification standard, or sample divided by the absorbance of the negative control * 100.

- Limitations

  - This procedure is a screening test.  Additional analysis is required to report the identification of a controlled substance or controlled substance metabolite.  Refer to the DWI Blood Chemistry Unit Technical Procedure for DWI Blood Chemistry Analysis.

  - Refer to the individual Immunalysis ELISA kit information for cross-reactivities.

  - The addition of sodium azide to blood affects the analysis and is not suitable for analysis with this procedure.  Interference with the enzyme process may cause false positive and false negative results.  Refer to 12.5.

  - The data obtained from the ELISA analyzer is displayed with 5 significant figures.  Any trailing digits that are zero are not printed by the analyzer software.

  - Immunalysis requires a %CV of 15 or less to accept a batch of assays.  Their %CV is based on more than 2 duplicates (usually 20 or 40).  The DWI Blood Chemistry Unit Technical Procedure adds 5% to this requirement to factor other sources of variability (i.e. pipetting, diluting, etc.).

- Safety

  - Refer to the Health and Safety Manual.

- References

  - Immunalysis Tecan Freedom EVO 75 Workstation Operating Manual V2.2, Tecan Schweiz AG, Switzerland, 2007.

  - Instructions for Use for Magellan, V 1.2, Tecan.

  - Instructions for Use for Sunrise Microplate Absorbance Reader, V2.0, Tecan.

  - Instructions for Use for Hydroflex, V2.2, Tecan.

  - Levine, Barry ed., Principles of Forensic Toxicology. 4th edition. AACC Press, 2017, 149-169.

  - Immunalysis ELISA kit inserts available from  info@immunalysis.com.

  - Moffat, Anthony C. ed. Clarke’s Analysis of Drugs and Poisons, Volume 1, 4th edition, Pharmaceutical Press, 2011.

- Records

  - ELISA Instrument Log

  - ELISA Maintenance Log

  - ELISA Monthly Maintenance Log

  - ELISA QC Summary

  - ELISA Case Summary

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
|  |  | Updated section 3. |
| 1/16/15 | 2 | Updated for minor corrections and clarifications: 4.6.1.1.4, 4.6.1.1.5, 7.23, 7.24.4, 7.27 and 12.6 |
| 9/15/2017 | 3 | Added assay for Amphetamine: 4.3.1, 4.5.4, 4.6.1.1.10, 4.6.2.2.10, 7.7.4, and 7.8.5. Changed wording in 7.7.1 and 7.8.1 |
| 4/10/2018 | 4 | Updated safety manual name in 11. |
| 6/6/2019 | 5 | Removed requirement for printer in 4.1.4 and 7.3. Updated reference for QCC in 4.6.1.6. and 4.2.6.7. Removed 4.2.6.8. Added assay for Fentanyl in 4.3.5, 4.5.4, 4.6.1.1.11, 4.6.2.2.11, 7.6.4, and 7.7.5. Updated Acid/Base wash in 6.4.2.2. Updated 7.4. Changed evaluation from Absorbance to %B/B0 in 7.2.3, 7.2.7, and 9.4. Updated Quality Control Check title in 7.24. Updated %CV requirements in 7.24.1. Updated evaluation criteria for Negative Standards in 7.23.1 and 7.27. Updated evaluation criteria for Calibration Standards in 7.24.2, and 7.27. Updated date format in 7.28. Updated evaluation criteria for Verification Standards in 7.24.3 and 7.27. Added a limitation in 10.5. Updated references in 12.5 and 12.7. Added Records 13.4 and 13.5. |
| 6/8/20 | 6 | 4.5. Commercial Reagents - revised wording of 4.5.4 4.6.1. ELISA Calibration Solution - revised wording at 4.6.1.1. and changed expiration at 4.6.1.4. 4.6.2. ELISA Verification Solution - revised wording at 4.6.2.2 and changed expiration at 4.6.2.5. 7.28.2 - removed 7.28.2.14 requirement to be included in ELISA QCC file 7.29. - revised required information and documents to be included in case files |
| 1/20/21 | 7 | 3. Abbreviations - removed 3.2 - 3.4 4.2. Materials - Added pipets to 4.2.5 4.3. Reference Material - Deleted carisoprodol (4.3.3.) and zolpidem (4.3.11) Updated standard dilutions at 4.3.1, 4.3.3 – 4.3.9. 4.5. Commercial Reagents - Deleted carisoprodol, zolpidem (4.5.4) and Zolpidem TMB substrate (4.5.4.4) 4.6.1. ELISA Calibration Solution - Deleted carisoprodol (4.6.1.1.2) and zolpidem (4.6.1.1.9) Updated volumes added to Calibration Solution at 4.6.1.1.2, and 4.6.1.1.4 – 4.6.1.1.9 4.6.2. ELISA Verification Solution - Deleted carisoprodol (4.6.2.2.2) and zolpidem (4.6.2.2.9) Updated volumes added to Verification Solution at 4.6.2.2.2 and 4.6.2.2.4 – 4.6.2.2.9 7.6. Calibration Standard - Deleted compounds and adjusted volumes added to Calibration solution concentrations at 7.6.4. 7.7. Verification Standard - Deleted compounds and adjusted volumes added to Verification solution concentrations at 7.7.5. 7. Procedure - removed Zolpidem reference from 7.11, 7.14, 7.19.4, 7.28.2.4, 7.29.2.5. Removed elevated criteria from 7.27. revised 7.28.1 and 7.28.2 to store files electronically. |

# 9: Technical Procedure for Solid Phase Extraction of Basic Drugs for GC-MS Analysis

- Purpose / Scope - The procedure may be used to isolate basic drugs from blood using solid phase extraction for analysis by gas chromatography - mass spectrometry in the DWI Blood Chemistry Unit of the CCBI Crime Laboratory.

- Definitions

  - Quality control check – Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

  - MeCl – Methylene chloride

  - IPA - Isopropanol

  - QCC – Quality control check

- Equipment, Materials and Reagents

  - Equipment

    - pH meter

    - TurboVap Evaporator

    - Mechanical pipettors and corresponding tips

    - Volumetric flasks, class A

    - Volumetric pipets, class A

  - Materials

    - Test tubes, 16 x 125mm and 13 x 100mm

    - Test tube caps or stoppers

    - Vortex Mixer

    - Centrifuge

    - UCT Clean Screen DAU Solid Phase Extraction Columns

    - Deionized water

    - Positive pressure manifold

    - Nitrogen, UHP

  - Reagents

    - Critical Reagents

      - Negative blood

      - BSTFA with 1% TMCS (N,O-bis(trimethylsilyl)trifluoroacetamide with 1% trimethylchlorosilane)

    - Reference materials

      - Prazepam

    - Commercial reagents

      - Methylene Chloride, Optima grade  HPLC grade, unstabilized

      - Ammonium Hydroxide, ACS grade

      - Ethyl Acetate, ACS grade

      - Glacial Acetic Acid, ACS grade

      - Methanol, ACS grade

      - Hexane, ACS grade

      - Isopropanol, ACS grade

      - Disodium Hydrogen Phosphate, Anhydrous, ACS grade

      - Sodium Dihydrogen Phosphate, Monohydrate, ACS grade

    - Prepared Reagents

      - Acetic Acid,  100 mM

        - Add  2.9 mL glacial acetic acid to 400 mL deionized water in a 500 mL volumetric flask.

        - Mix and dilute to 500 mL with deionized water.

        - Lot Number:  Eight digit format year/month/day/AceticAcid1.0M/initials of preparer.

        - Expiration:  six months

        - Storage:   closed container

        - QCC:  Tests acidic with pH or litmus paper

      - Methylene chloride (MeCl) : Isopropanol (IPA) 4:1 with 2 % Ammonium Hydroxide

        - Thoroughly mix isopropanol with ammonium hydroxide, add methylene chloride and mix well.

        - Lot number:    Eight digit format year/month/day/Amm4:1MeCl:IPA/initials of preparer.

        - Expiration:  end of day

        - Storage:  closed container

        - QCC:  basic to litmus or pH paper

      - 0.1 M Monobasic sodium phosphate, NaH2PO4

        - Dissolve 1.38 grams monobasic sodium phosphate monohydrate, in deionized water in a 100 mL volumetric flask.

        - Mix and dilute to 100 mL with deionized water.

        - Lot number:  Eight digit format year/month/day/0.1M_NaH2PO4/initials of preparer

        - Expiration:  one year

        - Storage:  closed container

        - QCC:    Tests acidic with pH or litmus paper

      - 0.1 M Dibasic sodium phosphate, Na2HPO4

        - Dissolve 1.42 grams dibasic sodium phosphate, anhydrous, in 80 mL deionized water in a 100 mL volumetric flask.

        - Mix and dilute to 100 mL with deionized water.

        - Lot number:  Eight digit format year/month/day/0.1M_Na2HPO4/initials of preparer.

        - Expiration:  one year

        - Storage:  closed container

        - QCC:  Tests basic with pH paper or litmus paper.

      - 100mM Phosphate Buffer, pH 6.0

        - Dissolve 0.85 g Na2HPO4 and 6.07 g NaH2PO4-H2O in 400 mL DI water and dilute to 500mL.

        - Adjust pH to 6.0 +/- 0.1 with 0.1M monobasic sodium phosphate (lowers pH) or 0.1M dibasic sodium phosphate (raises pH).

        - Lot number:  Eight digit format year/month/day/0.1MPhosBuffer/initials of preparer

        - Storage:  refrigerate in closed container

        - Expiration:  one month

        - QCC:  Record final pH

      - Prazepam Internal Standard Stock Solution, 100 μg/ml

        - Dilute 1.0 ml of a 1.0 mg/ml prazepam reference standard solution to 10 ml in a class A volumetric flask.  Mix.

        - Lot number:  Eight digit format year/month/day/100μg/mlPrazepamISStock/initials of preparer.

        - Storage:  freezer in closed container

        - Expiration:  three years

        - QCC:  N/A, refer to 4.3.4.8.

      - Prazepam Internal Standard Solution, 2μg/ml

        - Pipet 1.0 ml Prazepam Internal Standard Stock solution, 100 ng/ml, into a 50 ml class A volumetric flask.  Dilute to volume with methanol.

        - Lot number:  Eight digit format year/month/day/2μg/mlPrazepamIS /initials of preparer

        - Storage:  refrigerate in closed container

        - Expiration:  one year

        - QCC:  successful negative control extraction

- Quality Control

  - Positive control

    - Prazepam is added to each basic extraction sample as an internal standard.  For each basic extraction sample, the mass spectrum of the prazepam internal standard must have a positive comparison to prazepam reference material and the signal-to-noise ratio for the prazepam internal standard gas chromatographic peak must be 5:1 or greater.  The signal-to-noise ratio is defined as the response at the peak height of the internal standard ÷ the response at the baseline or valley immediately before the internal standard signal.

  - Negative Control

    - For each extraction batch of blood samples prepare a negative control as directed in Section 6 with 2.0 mL of negative blood.

    - The mass spectrum of the prazepam internal standard of the negative control must have a positive comparison to prazepam reference material and the signal-to-noise ratio for the prazepam internal standard gas chromatographic peak must be 5:1 or greater.  The signal-to-noise ratio is defined as the response at the peak height of the internal standard ÷ the response at the baseline or valley immediately before the internal standard signal.

    - The negative control must not contain any controlled substances, controlled substance metabolites or any substance being identified in the sample.  The negative control shall be subjected to the same post extraction techniques as any case samples in the batch.

- Procedure

  - Allow all solutions and samples to equilibrate to room temperature.

  - Ensure that all blood samples are homogenous by shaking and/or vortexing.

    - If a homogenous sample cannot be obtained, make a notation in the worksheet detailing the condition of the sample and its handling.

  - To 1 mL of 100 mM phosphate buffer add 100 μl of the prazepam internal standard solution, mix/vortex.

  - Add 2.0 mL of blood sample to be analyzed (case sample or negative blood for the negative control) using a mechanical pipettor labeled “blood” and add 3 mL of deionized water.

  - Mix/vortex and let stand for 5 minutes to lyse red blood cells.

  - Mix/vortex sample and centrifuge for 10 minutes at >2000 RPM

  - Decant liquid portion of the sample into 2 mL of 100 mM phosphate buffer solution.

  - Place the Clean Screen DAU Extraction columns into the positive pressure manifold column rack and place the column rack onto the support arms of the manifold.

  - Turn on the nitrogen to the manifold and select the corresponding switches on the front and rear of the manifold to allow nitrogen flow to the corresponding columns.

  - Adjust the two pressure regulators on the front of the manifold.  Use the low pressure regulator to load or aspirate a sample or liquid from the column at a flow rate of 1 to 2 ml per minute, 1 psi or less.  Use the high pressure regulator to dry the column with a pressure setting of 80 psi.

  - Condition the column with 3 ml methanol.

  - Rinse the column with 3 ml 4:1 MeCl:IPA with 2% NH4OH

  - Condition the column with 3 ml deionized water

  - Condition the column with 1 ml 100 mM phosphate buffer

  - Load sample onto column

  - Rinse column with 3 ml deionized water

  - Rinse column with 1 ml acetic acid,  100 mM

  - Rinse column with 3 ml methanol

  - Dry column for 5 minutes at 80 psi nitrogen

  - Collect basic drugs with 3 ml 4:1 MeCl:IPA with 2% NH4OH.

  - Evaporate the solvent from the collection test tube using the TurboVap LV Evaporator, refer to the DWI Blood Chemistry Unit Technical Procedure for Biotage TurboVap LV Evaporator.  Remove the tubes immediately upon reaching dryness.

  - Reconstitute with 50 µl ethyl acetate or perform silylation as directed in 6.22.1 through 6.22.4 for improved analysis by GC-MS.

    - Derivatize the extract in the collection tube or transfer to a vial using ethyl acetate and evaporate the ethyl acetate.  Discontinue evaporation immediately upon reaching dryness.

    - Derivatize by adding 50 μl of BSTFA with 1% TMCS to the vial or test tube and capping.

    - Mix and heat the vial or test tube for 30 minutes at 80°C.

    - Remove from the heat source and allow the vial or test tube to cool.

  - Analyze by GC-MS, refer to the DWI Blood Chemistry Unit Technical Procedure for Gas Chromatography – Mass Spectrometry.

  - Record the following in the case file

    - Lot number of internal standard

    - Lot number of negative blood

    - Lot number of UCT Clean Screen DAU Solid Phase Extraction Columns

    - If applicable, the lot number of BSTFA with 1 % TMCS

- Limitations

  - Refer to the references and other published chemical information as needed to determine the fraction in which a target analyte is expected to elute.  Typically barbiturates, carisoprodol, meprobamate and some benzodiazepines elute in the acidic/neutral fraction.  Typically alkaloids, amphetamines, opiates, zolpidem, tramadol and most benzodiazepines elute in the basic fraction.

  - Refer to the references and other published chemical information as needed to determine the need for derivatization.  Typically morphine and benzoylecgonine need to be derivatized for detection by GC-MS.  Some benzodiazepines and other substances may need to be derivatized for detection by GC-MS, e.g., a benzodiazepine or opiate ELISA indication and no corresponding substance detected in a non-derivatized sample.

  - Do not allow the solid phase extraction columns to dry during the extraction other than at steps indicated.

  - Store solid phase extraction columns in a closed container.

- Safety

  - Refer to the CCBI Health and Safety Manual

- References

  - UCT Solid Phase Extraction Manual, United Chemical Technologies Inc. Bristol, PA., (2011) 9 –12, 56 – 58.

  - BSTFA with 1 % TMCS Product Specification, Sigma-Aldrich Co, (1997).

  - Moffat, Anthony C. ed. Clarke’s Analysis of Drugs and Poisions, Volume 1, 4th edition, Pharmaceutical Press, 2011.

  - O’Neal, Maryadele J. ed. The Merck Index – An Encyclopedia of Chemicals, Drugs and Biologicals, Merck & Co Inc., Whitehouse Station, NJ, (2006).

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
| 11/8/13 | 2 | Update to line 7.1 due to incorporation of DBCTP14 |
| 1/16/15 | 3 | Updated for minor corrections and changes: 1, 4.3.4.2, 4.3.4.3.1, 4.3.4.4.1, 5.1, 5.2.2 and 6.22.1. |
| 9/21/15 | 4 | Require use of unstabilized HPLC grade methylene chloride |
| 8/11/16 | 5 | Require use of Optima grade methylene chloride |
| 2/14/17 | 6 | Updated 4.3.4.1, 4.3.4.1.1, and 6.17 to change Acetic Acid Concentration to 100 mM to be consistent with United Chemical Technologies published procedures. |
| 4/10/2018 | 7 | Updated safety manual name in 8. |

# 10: Technical Procedure for Solid Phase Extraction of Drugs for GC-MS Analysis

- Purpose / Scope – This procedure is used to isolate acidic / neutral drugs from blood, when an acidic / neutral fraction is collected, using solid phase extraction for analysis by gas chromatography – mass spectrometry in the DWI Blood Chemistry Unit of the CCBI Crime Laboratory.  This procedure may also be used to isolate basic drugs from blood when a basic fraction is collected.

- Definitions

  - Quality control check – Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

  - IPA – Isopropanol

  - MeCl – Methylene chloride

  - QCC – Quality control check

- Equipment, Materials and Reagents

  - Equipment

    - pH meter

    - TurboVap Evaporator

    - Mechanical pipettors and corresponding tips

    - Volumetric flasks, class A

    - Volumetric pipets, class A

  - Materials

    - Test tubes, 16 x 125mm and 13 x 100mm

    - Test tube caps or stoppers

    - Vortex Mixer

    - Centrifuge

    - UCT Clean Screen DAU Solid Phase Extraction Columns

    - Deionized water

    - Positive pressure manifold

    - Nitrogen, UHP

  - Reagents

    - Critical Reagents

      - Negative blood

      - BSTFA with 1% TMCS (N,O-bis(trimethylsilyl)trifluoroacetamide with 1% trimethylchlorosilane)

    - Reference materials

      - Hexobarbital

      - Prazepam

    - Commercial reagents

      - Methylene Chloride, Optima grade HPLC grade, unstabilized

      - Ammonium Hydroxide, ACS grade,

      - Ethyl Acetate, ACS grade

      - Glacial Acetic Acid, ACS grade

      - Methanol, ACS grade

      - Hexane, ACS grade

      - Isopropanol, ACS grade

      - Disodium Hydrogen Phosphate, Anhydrous, ACS grade

      - Sodium Dihydrogen Phosphate, Monohydrate, ACS grade

    - Prepared Reagents

      - Acetic Acid,  100 mM

        - Add  2.9 mL glacial acetic acid to 400 mL deionized water in a 500 mL volumetric flask.

        - Mix and dilute to 500 mL with deionized water.

        - Lot Number:  Eight digit format year/month/day/AceticAcid1.0M/initials of preparer.

        - Expiration:  six months

        - Storage:   closed container

        - QCC:  Tests acidic with pH or litmus paper

      - Methylene chloride (MeCl) : Isopropanol (IPA) 4:1 with 2 % Ammonium Hydroxide

        - Thoroughly mix isopropanol with ammonium hydroxide, add methylene chloride and mix well.

        - Lot number:    Eight digit format year/month/day/Amm4:1MeCl:IPA/initials of preparer.

        - Expiration:  end of day

        - Storage:  closed container

        - QCC:  basic to litmus or pH paper

      - 0.1 M Monobasic sodium phosphate, NaH2PO4

        - Dissolve 1.38 grams monobasic sodium phosphate, monohydrate, in deionized water in a 100 mL volumetric flask.

        - Mix and dilute to 100 mL with deionized water.

        - Lot number:  Eight digit format year/month/day/0.1M_NaH2PO4/initials of preparer

        - Expiration:  one year

        - Storage:  closed container

        - QCC:    Tests acidic with pH or litmus paper

      - 0.1 M Dibasic sodium phosphate, Na2HPO4

        - Dissolve 1.42 grams dibasic sodium phosphate, anhydrous, in 80 mL deionized water in a 100 mL volumetric flask.

        - Mix and dilute to 100 mL with deionized water.

        - Lot number:  Eight digit format year/month/day/0.1M_Na2HPO4/initials of preparer.

        - Expiration:  one year

        - Storage:  closed container

        - QCC:  Tests basic with pH paper or litmus paper.

      - 100mM Phosphate Buffer, pH 6.0

        - Dissolve 0.85 g Na2HPO4 and 6.07 g NaH2PO4-H2O in 400 mL DI water and dilute to 500mL.

        - Adjust pH to 6.0 +/- 0.1 with 0.1M monobasic sodium phosphate (lowers pH) or 0.1M dibasic sodium phosphate (raises pH).

        - Lot number:  Eight digit format year/month/day/0.1MPhosBuffer/initials of preparer

        - Storage:  refrigerate in closed container

        - Expiration:  one month

        - QCC:  Record final pH

      - Hexobarbital Internal Standard. 40 μg/ml

        - Dilute 1 ml of a 1.0 mg/mL hexobarbital reference material solution to volume in a 25 ml class A volumetric flask with methanol.  Mix.

        - Lot number:  Eight digit format year/month/day/40μg/mlHexIS/initials of preparer.

        - Storage:  refrigerator

        - Expiration:  one year

        - QCC:  successful negative control extraction

      - Prazepam Internal Standard Stock Solution, 100 μg/ml

        - Dilute 1.0 ml of a 1.0 mg/ml prazepam reference standard solution to 10 ml in a class A volumetric flask.  Mix.

        - Lot number:  Eight digit format year/month/day/100μg/mlPrazepamISStock/initials of preparer.

        - Storage:  freezer in closed container

        - Expiration:  three years

        - QCC:  N/A, refer to 4.3.4.8.

      - Prazepam Internal Standard Solution, 2μg/ml

        - Pipet 1.0 ml Prazepam Internal Standard Stock solution, 100 ng/ml, into a 50 ml class calss A volumetric flask.  Dilute to volume with methanol.

        - Lot number:  Eight digit format year/month/day/2μg/mlPrazepamIS /initials of preparer

        - Storage:  refrigerate in closed container

        - Expiration:  one year

        - QCC:  successful negative control extraction

- Quality Control

  - Positive control

    - Hexobarbital is added to each acidic / neutral extraction sample as an internal standard.  For each acidic / neutral extraction sample, the mass spectrum of the hexobarbital internal standard must have a positive comparison to hexobarbital reference material and the signal-to-noise ratio for the hexobarbital internal standard gas chromatographic peak must be 5 2:1 or greater.  The signal-to-noise ratio is defined as the response at the peak height of the internal standard ÷ the response at the baseline or valley immediately before the internal standard signal.

    - Prazepam is added to each basic extraction sample as an internal standard.  For each basic extraction sample, the mass spectrum of the prazepam internal standard must have a positive comparison to prazepam reference material and the signal-to-noise ratio for the prazepam internal standard gas chromatographic peak must be 5:1 or greater.  The signal-to-noise ratio is defined as the response at the peak height of the internal standard ÷ the response at the baseline or valley immediately before the internal standard signal.

  - Negative Control

    - For each extraction batch of blood samples prepare a negative control as directed in Section 6 with 2.0 mL of negative blood.

    - The mass spectrum of the hexobarbital internal standard of the negative control must have a positive comparison to hexobarbital reference material and the signal-to-noise ratio for the hexobarbital internal standard gas chromatographic peak must be 5:1 or greater.  The signal-to-noise ratio is defined as the response at the peak height of the internal standard ÷ the response at the baseline or valley immediately before the internal standard signal.

      - For each basic fraction the mass spectrum of the prazepam internal standard of the negative control must have a positive comparison to prazepam reference material and the signal-to-noise ratio for the prazepam internal standard gas chromatographic peak must be 5 2:1 or greater.  The signal-to-noise ratio is defined as the response at the peak height of the internal standard ÷ the response at the baseline or valley immediately before the internal standard signal.

    - The negative control must not contain any controlled substances, controlled substance metabolites or any substance being identified in the sample.  The negative control shall be subjected to the same post extraction techniques as any case samples in the batch.

- Procedure

  - Allow all solutions and samples to equilibrate to room temperature.

  - Ensure that all blood samples are homogenous by shaking and/or vortexing.

    - If a homogenous sample cannot be obtained, make a notation in the worksheet detailing the condition of the sample and its handling.

  - To 1 mL of 100 mM phosphate buffer add 50 μl of the hexobarbital internal standard solution.  If a basic fraction is to be collected, add 100 μl of the prazepam internal standard solution.   Mix/vortex.

  - Add 2.0 mL of blood sample to be analyzed (case sample or negative blood for the negative control) using a mechanical pipettor labeled “blood” and add 3 mL of deionized water.

  - Mix/vortex and let stand for 5 minutes to lyse red blood cells.

  - Mix/vortex sample and centrifuge for 10 minutes at >2000 RPM

  - Decant liquid portion of the sample into 2 mL of 100 mM phosphate buffer solution.

  - Place the Clean Screen DAU Extraction columns into the positive pressure manifold column rack and place the column rack onto the support arms of the manifold.

  - Turn on the nitrogen to the manifold and select the corresponding switches on the front and rear of the manifold to allow nitrogen flow to the corresponding columns.

  - Adjust the two pressure regulators on the front of the manifold.  Use the low pressure regulator to load or aspirate a sample or liquid from the column at a flow rate of 1 to 2 ml per minute, 1 psi or less.  Use the high pressure regulator to dry the column with a pressure setting of 80 psi.

  - Condition the column with 3 ml methanol.

  - Rinse the column with 3 ml 4:1 MeCl:IPA with 2% NH4OH

  - Condition the column with 3 ml deionized water

  - Condition the column with 1 ml 100 mM phosphate buffer

  - Load sample onto column

  - Rinse column with 3 ml deionized water

  - Rinse column with 1 ml acetic acid, 100 mM

  - Dry column for 5 minutes with nitrogen

  - Wash the column with 2 mL of hexane

  - Collect acidic / neutral drugs with 6 ml methylene chloride

    - To also collect a basic fraction:

      - Rinse column with 3 ml methanol

      - Dry column for 2 minutes at 80 psi nitrogen

      - Collect basic drugs with 3 ml 4:1 MeCl:IPA with 2% NH4OH in a separate collection tube

  - Evaporate the solvent from the collection test tube using the TurboVap LV Evaporator, refer to the DWI Blood Chemistry Unit Technical Procedure for Biotage TurboVap LV Evaporator.  Remove the tubes immediately upon reaching dryness.

  - Reconstitute acidic / neutral fractions with 100 µl ethyl acetate or perform silylation as directed in 6.22.1. through 6.22.4. for improved analysis by GC-MS.  Acidic / neutral fractions may be diluted with additional ethyl acetate to achieve improved chromatography.  Do not dilute the negative control fraction with more ethyl acetate than used for the sample fraction.  Record the amount of ethyl acetate used.  Reconstitute basic fractions with 50 µl ethyl acetate or perform silylation as directed in 6.22.1 through 6.22.4 for improved analysis by GC-MS.

    - Derivatize the extract in the collection tube or transfer to a vial using ethyl acetate and evaporate the ethyl acetate.  Discontinue evaporation immediately upon reaching dryness.

    - Derivatize by adding 50 μl of BSTFA with 1% TMCS to the vial or test tube and capping.

    - Mix and heat the vial or test tube for 30 minutes at 80°C.

    - Remove from the heat source and allow the vial or test tube to cool.

  - Analyze by GC-MS, refer to the DWI Blood Chemistry Unit Technical Procedure for Gas Chromatography – Mass Spectrometry.

  - Record the following in the case file

    - Lot number of internal standard(s)

    - Lot number of negative blood

    - Lot number of UCT Clean Screen DAU Solid Phase Extraction Columns

    - If applicable, the lot number of BSTFA with 1 % TMCS

- Limitations

  - Refer to the references and other published chemical information as needed to determine the fraction in which a target analyte is expected to elute.  Typically barbiturates, carisoprodol, meprobamate and some benzodiazepines elute in the acidic/neutral fraction.  Typically alkaloids, amphetamines, opiates, zolpidem, tramadol and most benzodiazepines elute in the basic fraction.

  - Refer to the references and other published chemical information as needed to determine the need for derivatization.  Typically morphine and benzoylecgonine need to be derivatized for detection by GC-MS.  Some benzodiazepines and other substances may need to be derivatized for detection by GC-MS, e.g., an opiate or benzodiazepine ELISA indication and no corresponding substance detected in a non-derivatized sample.

  - Do not allow the solid phase extraction columns to dry during the extraction other than at steps indicated.

  - Store solid phase extraction columns in a closed container.

- Safety

  - Refer to the CCBI Health and Safety Manual

- References

  - UCT Solid Phase Extraction Manual, United Chemical Technologies Inc. Bristol, PA., (2011) 9 –12, 56 – 58.

  - BSTFA with 1 % TMCS Product Specification, Sigma-Aldrich Co, (1997).

  - Moffat, Anthony C. ed. Clarke’s Analysis of Drugs and Poisions, Volume 1, 4th edition, Pharmaceutical Press, 2011.

  - O’Neal, Maryadele J. ed. The Merck Index – An Encyclopedia of Chemicals, Drugs and Biologicals, Merck & Co Inc., Whitehouse Station, NJ, (2006).

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
| 11/8/13 | 2 | Incorporation of DBCTP14 |
| 1/16/15 | 3 | Updated for minor corrections and additions: 1, 4.3.4.2, 4.3.4.3.1, 4.3.4.4.1, 4.3.4.7.1, 4.3.4.9.1, 5.1.1, 5.1.2, 5.2.2, 6.3 and 6.22.1. |
| 9/21/15 | 4 | Require use of unstabilized HPLC grade methylene chloride |
| 8/11/16 | 5 | Require use of Optima grade methylene chloride |
| 2/14/17 | 6 | Updated 4.3.4.1, 4.3.4.1.1, and 6.17 to change Acetic Acid Concentration to 100 mM to be consistent with United Chemical Technologies published procedures. |
| 4/10/2018 | 7 | Updated Title. Updated safety manual name in 8. |

# 11: Technical Procedure for Supported Liquid Extraction (SLE) of THC, THC-OH and THC-COOH for LCMS-MS Analysis

- Purpose / Scope – This procedure is used to isolate delta-9-tetrahydrocannabinol, 11-nor-9-hydroxy-delta-9-Tetrahydrocannabinol, and 11-nor-9-carboxy-delta-9-tetrahydrocannabinol from blood using a ToxBox custom analytical plate and a Biotage 48-well SLE plate for analysis by Liquid chromatography - mass spectrometry in the DWI Blood Chemistry Unit of the CCBI Crime Laboratory.

- Definitions

  - Quality control check (QCC) - Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

- Equipment, Materials and Reagents

  - Equipment

    - Biotage SPE Dry Evaporator

    - Positive Pressure Manifold

    - Shaking Incubator

    - Mechanical pipettors and corresponding tips

    - Graduated Cylinder

  - Materials

    - Vortex mixer

    - Biotage 48 well Isolute SLE+ 1mL Supported Liquid Extraction Plate

    - Deionized water

    - Nitrogen, UHP

  - Reagents

    - Critical Reagents

      - Negative blood

    - Reference materials

      - ToxBox THC custom analytical plate containing:

        - delta-9-tetrahydrocannabinol (THC)

        - 11-nor-9-hydroxy-delta-9-tetrahydrocannabinol (THC-OH)

        - 11-nor-9-carboxy-delta-9-tetrahydrocannabinol (THC-COOH)

        - delta-9-tetrahydrocannabinol-D3 (THC-D3)

        - 11-nor-9-hydroxy-delta-9-tetrahydrocannabinol-D3 (THC-OH-D3)

        - 11-nor-9-carboxy-delta-9-tetrahydrocannabinol-D9 (THC-COOH-D9)

      - Delta-9-tetrahydrocannabinol (THC)

      - 11-nor-9-hydroxy-delta-9-tetrahydrocannabinol (THC-OH)

      - 11-nor-9-carboxy-delta-9-tetrahydrocannabinol (THC-COOH)

    - Commercial reagents

      - Formic Acid, LC/MS grade

      - Tert-Butyl Methyl Ether (MTBE), LC/MS grade

      - Methanol, LC/MS grade

      - Acetonitrile, LC/MS grade

      - Hexane, LC/MS grade

    - Prepared Reagents

      - 0.1% Formic Acid Solution

        - Mix deionized water and Formic Acid in a ratio of 1000:1.

        - Lot Number: Eight digit format year/month/day/0.1%Formicacid/initials of preparer.

        - Expiration:  1 month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - LCMS-MS Mobile Phase A - 0.1% Formic Acid in Water

        - Mix deionized water and Formic acid in a ratio of 1000:1.

        - Lot Number:  Eight digit format year/month/day/MPA0.1%FAH2O/initials of preparer.

        - Expiration:  1 month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - Screening LCMS-MS Mobile Phase B - 0.1% Formic Acid in Methanol

        - Mix Methanol and Formic acid in a ratio of 1000:1.

        - Lot Number:  Eight digit format year/month/day/MPB0.1%FAMeOH/initials of preparer.

        - Expiration:  one month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - Quantitative LCMS-MS Mobile Phase B - 0.1% Formic Acid in Acetonitrile

        - Mix Acetonitrile and Formic Acid in a ratio of 1000:1.

        - Lot Number:  Eight digit format year/month/day/MPB0.1%FACH3CN/initials of preparer.

        - Expiration:  one month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - Cannabinoid Verification Solution

        - Prepare this solution with reference materials from a different supplier or different supplier lot than those used in the ToxBox THC custom analytical plates.

        - Prepare a solution that contains 1 μg/ml of THC, THC-COOH, and THC-OH.  Example: Pipet 1.0 ml of a 100 μg/ml THC reference material solution,  1.0 ml of a 100 μg/ml THC-COOH reference material solution, and 1.0  of a 100 μg/ml THC-COOH reference material solution to a 100 ml volumetric flask and dilute to volume with methanol.

        - Lot Number:  Eight digit format year/month/day/CanVer/initials of preparer.

        - Expiration:  one year

        - Storage:   refrigerator

        - QCC:  Successful QC positive control

      - 50 ng/ml Cannabinoid Verifier

        - Add 100 µL of the Cannabinoid Verification Solution to 1.9 mL of negative blood.

        - Lot Number: eight digit formal year/month/day/50ng/mlCanVer/initals of preparer

        - Expiration: end of the day

        - Storage: closed container

        - QCC: Successful QC positive control

- Calibration

  - Calibrate the instrument each time that an extraction batch is analyzed and after instrument maintenance that may affect the calibration as determined by the DWI Blood Chemistry Technical Leader.

  - Calibration Standards

    - Screening plate THC and THC-OH:

Level 0 - Negative Calibration Standard

Level 1 - 1 ng/ml of THC and THC-OH

Level 2 - 5 ng/ml of THC and THC-OH

Level 3 - 25 ng/ml of THC and THC-OH

Level 4 - 100 ng/ml of THC and THC-OH

The QC levels are the following:

QC1 - 2 ng/ml of THC and THC-OH

The internal standard concentration is 100ng/ml of each THC-d3, THC-OH-d3, THC-COOH-d9.

    - Screening plate THC-COOH:

Level 0 - Negative Calibration Standard

Level 1 - 5 ng/ml of THC-COOH

Level 2 - 25 ng/ml of THC-COOH

Level 3 - 100 ng/ml of THC-COOH

Level 4 - 250 ng/ml of THC-COOH

The QC levels are the following:

QC1 - 5 ng/ml of THC-CO-OH

The internal standard concentration is 100ng/ml of each THC-d3, THC-OH-d3, THC-COOH-d9.

    - Quantitation plate THC and THC-OH:

Level 0 - Negative Calibration Standard

Level 1 - 1 ng/ml of THC and THC-OH

Level 2 - 5 ng/ml of THC and THC-OH

Level 3 - 10 ng/ml of THC and THC-OH

Level 4 - 25 ng/ml of THC and THC-OH

Level 5 - 50 ng/ml of THC and THC-OH

Level 6 - 75 ng/ml of THC and THC-OH

Level 7 - 100 ng/ml of THC and THC-OH

The QC levels are the following:

QC Low - 2 ng/ml of THC and THC-OH

QC Mid - 25 ng/ml of THC and THC-OH

QC High - 75 ng/ml of THC and THC-OH

The internal standard concentration is 100ng/ml of each THC-d3, THC-OH-d3, THC-COOH-d9.

    - Quantitation plate THC-COOH:

Level 0 - Negative Calibration Standard

Level 1 - 5 ng/ml of THC-COOH

Level 2 - 10 ng/ml of THC-COOH

Level 3 - 25 ng/ml of THC-COOH

Level 4 - 50 ng/ml of THC-COOH

Level 5 - 75 ng/ml of THC-COOH

Level 6 - 100 ng/ml of THC-COOH

Level 7 - 250 ng/ml of THC-COOH

The QC levels are the following:

QC Low - 5 ng/ml of THC-COOH

QC Mid - 50 ng/ml of THC-COOH

QC High - 100 ng/ml of THC-COOH

The internal standard concentration is 100ng/ml of each THC-d3, THC-OH-d3, THC-COOH-d9.

- Procedure

  - Remove analytical plate, blood and samples from cold storage and allow to reach room temperature.

  - Ensure that all blood samples are homogenous by shaking and/or vortexing.

    - If a homogenous sample cannot be obtained, make a notation in the worksheet detailing the condition of the sample and its handling.

  - Add 1.0 mL negative blood to wells of analytical plate containing standards/QCs.  Add 1.0 mL of the Cannabinoid Verifier to one well containing only internal standard and 1.0 mL of case sample blood to remaining wells containing only internal standard.   Ensure that the blood samples are added consistent with the plate map.

  - Place plate on the shaking incubator at approximately 900 rpm for 15 minutes.

  - Add 0.5 mL 0.1% formic acid solution to all wells.  Mix several times via aspiration/dispensing.

  - Place plate on the shaking incubator at approximately 900 rpm for 15 minutes.

  - Transfer 800 µL of the blood + formic acid mixture to corresponding wells of a 48-well Biotage SLE plate.

  - Apply positive pressure until solution penetrates well sorbent approximately 50% (or until no liquid remains on top of sorbent) - approximately 4 seconds.

  - Wait 5 minutes for the sample to completely absorb.

  - Add 2.25 mL MTBE to all wells and allow to flow for 5 minutes under gravity.

  - Apply positive pressure to complete elution – approximately 20 seconds.

  - Add 2.25 mL hexane and allow to flow for 5 minutes under gravity.

  - Apply positive pressure to complete elution – approximately 20 seconds.

  - Evaporate to dryness in the SPE Dry at approximately 35°C under a constant flow of nitrogen.  Temperatures for this step are not critical and it is appropriate to use equipment temperature readings.

  - Reconstitute each sample with 100 µL of methanol.  Transfer samples to a clean 96 well plate for LCMS-MS analysis.  Heat seal foil on top to prevent evaporation.

  - Analyze all samples immediately or refrigerate for future analysis.

- Limitations

- Safety

  - Refer to the CCBI Health and Safety Manual

- References

  - Extraction of THC, Hydroxy-THC and Carboxy-THC from Whole Blood Using ISOLUTE SLE+ Application Note 2015, Biotage. AN840.

  - Biotage Pressure+ 96 User Manual, Biotage, Uppsala, Sweden.

  - Isotemp Shake Touch Instruction Manual,

Fisher Scientific, Revision 8, 2/2018.

- Records

  - Cannabinoid QCC file

  - Case record

  - Reagent Log

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
|  |  | Updated section 3. |
| 1/16/15 | 2 | Corrected reference to technical leader |
| 4/10/18 | 3 | Updated safety manual name in 9. |
| 10/28/19 | 4 | Revised entire procedure to SLE extraction of THC, THC-OH, THC-COOH for LCMS-MS Analysis. |
| 01/24/20 | 5 | 3. Abbreviations - removed abbreviations 4.3.2. Reference Material - added THC, THC-OH, and THC-COOH reference material. 4.3.4 Prepared Reagents - added Screening Mobile Phase B, Cannabinoid Verification Solution, and 50ng/ml Cannabinoid Verifier 5.3 through 6. - moved to DBCTP15 6. Procedure - added addition of cannabinoid Verifier to plate preparation |

# 12: Technical Procedure for Gas Chromatography/Mass Spectrometry (GC-MS)

- Purpose / Scope - This procedure provides direction for the initial setup, performance checks and usage of gas chromatograph – mass spectrometer instruments in the DWI Blood Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification Crime Laboratory.

- Definitions

  - Performance verification - The initial confirmation of the reliability of a previously or externally validated method or instrument.

  - Quality control check - Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

  - Reference Material – Material sufficiently homogeneous and stable with reference to specified properties, which has been established to be fit for its intended use in measurement or in examination of nominal properties

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

- Equipment, Materials and Reagents

  - Equipment

    - Agilent 7890 gas chromatograph with Agilent 5975 series mass selective detector with Agilent automatic liquid sampler and tray

    - Computer with Agilent Analytical MSD Productivity ChemStation software and printer

  - Reference Materials

    - Multi-component drug solution containing alprazolam, clonazepam, diazepam, flunitrazepam, lorazepam, nitrazepam, oxazepam and temazepam

    - Prazepam

    - Hexobarbital

  - Materials

    - Sample vials, caps and inserts

    - ALS Syringe, 10μl straight, fixed needle, 23/42/HP

    - DB5-MS column, 12 m X 0.200 mm X 0.33 µm for 70TOX1 and 70TOX1DERIV methods

    - Agilent inlet liner, splitless, single taper with glass wool, deactivated, 900 µl, 4mm ID

    - Agilent liner O-ring

    - Merlin microseal

    - Septum wrench

    - Tweezers

    - Clean, lint free, non-nylon gloves

    - Wrenches, ¼ inch and ½ inch

    - Gold plated inlet seal with cross and 0.375 outer diameter washer

    - Star or Torx screwdriver

    - Flat head screwdriver, large

    - Hex key, 5 mm

    - Inland 45 pump oil

    - Funnel

    - Hex ball driver, 1.5 mm

    - Hex ball driver, 2.0 mm

    - Wrench, open-end, 10 mm

    - Alumina abrasive powder

    - Cotton swabs

    - Ultrasonic bath

  - Commercial Reagents

    - Methanol, Optima grade

    - Ethyl acetate, Optima grade

    - Helium gas, Grade 5.0

    - Perfluorotributylamine [PFTBA], neat

  - Prepared Reagents

    - Monthly QCC solution

      - Mix hexobarbital reference material, prazepam reference material and benzodiazpine  reference material solution from 4.2.1 with ethyl acetate to yield a solution containing 10 ng/μl of each component.

      - Lot Number:  Eight digit format year/month/day/BenzoMix/initials of preparer.

      - Expiration:  one year

      - Storage:   closed container in freezer

      - QCC:  Successful monthly QCC

- Standards and Controls

  - An electronic GC-MS logbook shall be maintained.  The logbook shall contain the GC-MS Activity Log, GC-MS Daily QCC Log, the GC-MS Monthly QCC Log, the GC-MS Maintenance Log and any manufacturer’s certificates, performance verification documentation, QCC printouts and maintenance documentation.

    - GC-MS Activity Log

      - Record the date, sequence identification, initials of operator, and any comments for each sequence analyzed on the DWI Blood Chemistry GC-MS Activity Log.

      - Upon completion of each sequence, maintain the sequence in the logbook.

      - Record any error messages on the DWI Blood Chemistry GC-MS Activity Log.

    - GC-MS Daily QCC log

      - Record all Daily QCC’s on the Daily QCC log with the date, initials, and results and any comments.

      - Place Daily QCC printouts in the instrument logbook.

    - Monthly QCC log

      - Save all Monthly QCC’s on the Monthly QCC log with the date, initials, results and any comments, refer to 5.7.

      - Save Monthly QCC printouts in the instrument logbook.

  - When the GC-MS has been placed out of service for maintenance, malfunction or leaving direct control of the Laboratory the DWI Blood Chemistry Technical Leader shall evaluate the instrument and determine if any additional quality control checks are needed to ensure instrument performance.  At a minimum, a Daily QCC must be successfully performed prior to placing the instrument back in service, refer to 5.6.

    - If maintenance is performed that may affect retention times, a monthly QCC, refer to 5.7., shall be performed before the instrument is placed back in service.

  - The Blood Drug Chemist shall record any malfunctions or error messages in the GC-MS Activity Log, notify the DWI Blood Chemistry Technical Leader of any malfunctions or error messages and place the instrument out of service by marking the GC-MS Activity Log “Out of Service.”

  - The DWI Blood Chemistry Technical Leader shall correct any problems with the instrument or request service.  The DWI Blood Chemistry Technical Leader shall examine the effect(s), if any, of a malfunction or error message on analysis results and implement the CCBI Laboratory Procedure for Corrective and Preventive Action as required.

  - Negative Quality Control Check

    - Negative QCC’s are performed prior to each sample injection, refer to 7.3.

  - Daily Quality Control Check – AutoTune

    - Perform an Autotune (atune) with Perfluorotributylamine (PFTBA) as the tuning standard prior to beginning the first sample sequence each day the instrument is in use.

      - Sample sequences that continue overnight may be allowed to complete without performing a new tune provided that they do not extend more than 36 hours beyond the time of the tune.

    - Compare the atune report to previous ones.  Record any major variations on the Daily QCC log and notify the DWI Blood Chemistry Technical Leader.

    - The mass assignments of the tuning masses, 69.00, 219.00, and 502.00, in the upper part of the tune report must be within ± 0.2 amu.

    - The peak widths on the tune report for masses 69.00, 219.00, and 502.00, must be within +/- 0.10 amu of 0.55 (i.e., 0.45-0.65) amu at 50% peak height (PW50) and the peaks should be generally be smooth and symmetrical.

    - The base peak on the tune report must be identified as mass 69 or 219.  The relative abundance ratio of mass 219 to mass 69 must be > 40% and the relative abundance ratio of mass 502 to mass 69 must be > 2.4%.

    - The 70/69 isotopic ratio must be from 0.5 – 1.6, the 220/219 ratio must be from 3.2 – 5.4, and the 503/502 the ratio must be from 7.9 – 12.3.

    - The abundance of any peaks less than 69 amu should not be greater than 10 % of the abundance of the base peak.

      - Peaks at 18, 28 or 32 amu are indicative of water, nitrogen and oxygen, respectively, and may indicate an air leak.  Other peaks may indicate gas impurities.

      - If an air leak is detected, isolate the leak and tighten fittings to correct the leak and perform another atune.

      - If the problem persists, place the instrument out of service by marking the activity log “out of service” and notify the DWI Blood Chemistry Technical Leader.

The DWI Blood Chemistry Technical Leader shall correct any problems with the instrument or request service.

    - If any parameter listed in 5.6.2. - 5.6.7. does not meet the requirements listed, document the deviation on the atune report and on the GC-MS daily QCC log.

      - Perform another atune.

      - If the problem persists, place the instrument out of service by marking the activity log “out of service” and notify the DWI Blood Chemistry Technical Leader.

      - The DWI Blood Chemistry Technical Leader shall correct any problems with the instrument or request service.

      - The daily QCC must be successfully completed prior to placing the instrument back in service.

    - Save the atune report in the instrument logbook and record the results in the GC-MS daily QCC log.

  - Monthly Quality Control Check

    - The multi-component drug solution from 4.5.1 shall be injected on the method “70TOX1,” refer to 7.1., each month the instrument is in use to verify instrument performance.

    - The solution shall, when feasible, be run during the first seven calendar days of each month.

      - If the standard solution is not run during the first seven calendar days of the month, the instrument shall be out of service until the standard solution is successfully run.

    - Name the multi-component reference material solution data files with “MC” followed by the numerical year and month designation.  Name the corresponding solvent blank with the same designation followed by “-b”, indicating that it is a blank.

    - Visually examine the TIC of the monthly QCC solution for chromatographic quality and resolution.  All components must exhibit visually symmetrical peaks that are visually baseline resolved or for peaks separated by 0.2 minutes or less, visually resolved at half height.  Abundances should be comparable to those of the previous monthly QCC.

    - Perform a GC relative retention time comparison, refer to 7.7.2., for the oxazepam, temazepam and alprazolam components of the monthly QCC solution.

    - Perform a mass spectral comparison, refer to 7.7.1., for the oxazepam, temazepam and alprazolam components of the monthly QCC solution to reference material.

    - Each component specified must have a positive comparison for 5.7.5. and 5.7.6.

      - Record any component that does have a positive comparison on the instrument monthly QCC log, place the instrument out of service by marking the GC-MS activity log “out of service” and notify the DWI Blood Chemistry Technical Leader.

      - The DWI Blood Chemistry Technical Leader shall correct any problems with the instrument or request service.  The DWI Blood Chemistry Technical Leader shall examine the effect(s), if any, on analysis results and implement the CCBI Laboratory Procedure for Corrective and Preventive Action as required.

      - The monthly QCC must be successfully completed prior to placing the instrument back in service.

    - Save the monthly QCC TIC with the retention times displayed, the mass spectrum of oxazepam, temazepam, and alprazolam and the corresponding blank TIC.

      - Record the lot number of the solution on the TIC.

      - Save the generated data in the instrument logbook folder.

      - Record the monthly check in the monthly QCC log.

  - Performance Verification for New Instrumentation

    - New GC-MS instruments shall be installed by a manufacturer representative or approved vendor according to the manufacturer’s guidelines.

    - Prior to use, perform daily QCC’s, refer to 5.6., on three separate days.  The daily QCC’s must meet all specified requirements.

    - Prior to use, analyze the multi-component drug solution from 4.2.1 on each method (refer to 7.1.) on three separate days.

      - The mass spectra of each component must have a positive mass spectral comparison to reference material, refer to 7.7.1.

      - The relative retention times of each component must have a positive GC retention time comparison to reference material, refer to 7.7.2.

    - Label the instrument data with the lot number of the reference material and date.  Record the performance verification in the GC-MS logbook and place the data in the GC-MS logbook folder.

    - The performance verification must be reviewed and approved by the DWI Blood Chemistry Technical Leader prior to the instrument being used for casework.  The DWI Blood Chemistry Technical Leader shall record the review and approval in the GC-MS logbook.

- Maintenance

  - Place the instrument out of service prior to performing any column or mass spectrometer maintenance, by marking the GC-MS activity log “Out of Service.”  When the instrument is ready to be returned to service, mark the GC-MS activity log “Back in Service.”

  - Record all maintenance on the GC-MS maintenance log at the time it is performed with the name of the person performing the maintenance or repairs, the initials of the Blood Drug Chemist recording the maintenance or repairs, the date, the time, a description of the maintenance or repairs and a list of any parts replaced, the type of post maintenance QCC performed, and any additional comments.

  - Record the length of column trimmed in the activity log and a maintenance log.  If the column is trimmed, the instrument shall be out of service until a monthly QCC is successfully completed, refer to 5.7.

    - The DWI Blood Chemistry Technical Leader shall update the instrument log when the instrument is ready to be used for casework and save any generated data in the instrument logbook folder.

  - Suggested Routine maintenance – The routine maintenance schedule is a suggested guideline.  The maintenance schedule will be determined by the DWI Blood Chemistry Technical Leader based upon instrument usage and performance.

    - Wash Vials

      - Rinse and fill daily when in use.  Use methanol in the first wash vial and ethyl acetate in the second wash vial.

      - Required post-maintenance check:  None.

    - Merlin microseal

      - Replace yearly at a minimum when in use.

        - Press [Oven] and set the oven to 35°C. When the temperature reaches setpoint, turn the oven off.  Press [Front Inlet] and turn off the inlet temperature and pressure.

        - Be careful - The inlet fittings may be hot enough to cause burns.  Remove the septum retainer nut, using the wrench if the nut is hot or sticks.

        - Remove the old Merlin microseal, using tweezers if necessary.  Be sure that it is completely removed and take care to avoid gouging or scratching the interior of the septum head.

        - Press a new Merlin microseal into place firmly.

        - Replace the septum or Merlin microseal retainer nut, tightening it finger-tight until the c-ring is about 1 mm above the nut.  Avoid overtightening.

        - Using ChemStation load a method to return the GC to appropriate settings.  If prompted, do not save any method changes.

        - Allow the GC to return to the setpoints.

      - Record on the GCMS Maintenance Log.

      - Required post-maintenance check:  Successful daily QCC, refer to 5.6.

    - Syringe

      - Inspect monthly for cleanliness and ease of movement.  Replace as needed.

        - Mount the injector on a parking post.

        - Open the injector door.

        - Slide the syringe carriage to the top position.

        - Completely loosen the plunger thumb screw until it reaches the stop, and lift the plunger carrier off of the syringe plunger.

        - Open the syringe latch by swinging it in a counterclockwise direction.

        - Carefully pull the top of the syringe out of the flange guide, then lift the needle out of the needle support foot.

        - Carefully pass the new syringe needle through the guide hole in the needle support foot.

        - Align the syringe flange with the flange guide and press the syringe into place, keeping the needle end in the guide hole of the needle support foot. Make sure that the flat edge of the syringe flange faces out.

        - Close the syringe latch by swinging it clockwise until it snaps in place.

        - Slide the plunger carrier down until it is completely over the syringe plunger, and tighten the plunger thumb screw until finger- tight.

        - Manually move the plunger carrier up and down. If the syringe plunger does not move along with the carrier, repeat the previous steps until installed correctly.  Be sure the plunger thumb screw is secure and tight.  Verify that the needle is inside the guide hole of the needle support foot.  The needle should be straight and pass freely through the needle guide hole.  If the needle is bent or is outside the guide hole, remove the syringe and reinstall.

        - Close the injector door.

        - Mount the injector on the inlet.

      - Record on the GCMS Maintenance Log.

      - Required post-maintenance check:  None.

    - Liner

      - Replace monthly, at a minimum, when in use.

        - Press [Oven] and set the oven to 35°C. When the temperature reaches setpoint, turn the oven off.  Press [Front Inlet] and turn off the inlet temperature and pressure.

        - Be careful - The inlet fittings may be hot enough to cause burns.  Flip the inlet open.

        - Remove liner with tweezers, being careful not to break the liner.

        - Hold the new liner with tweezers or lint free tissue and place the o-ring on the liner about 2 to 3 mm from its top end.

        - Insert the liner straight down into the inlet and press gently to ensure it is seated.

        - Replace the inlet cover and flip the top into place.

        - Using ChemStation load a method to return the GC to appropriate settings.  If prompted, do not save any method changes.

        - Allow the GC to return to the setpoints.

      - Record on the GCMS Maintenance Log.

      - Required post-maintenance check:  Successful daily QCC, refer to 5.6.

    - Pump Oil

      - Change every six months.

        - Vent the MSD by selecting the vent option in Instrument Control of Chemstation.  Allow the vent cycle to run, when the cycle is complete and the temperatures are below 100 degrees Celsius turn off the MSD.

        - Press [Oven] and set the oven to 35°C.  Press [Front Inlet] and turn off the inlet temperature and pressure.  When the temperature reaches the setpoint turn the GC off.

        - Place a book or other object approximately two inches thick under the pump motor to tilt it up slightly.

        - Remove the fill cap.

        - Place a container under the drain plug.

        - Remove the drain plug.  Allow the pump oil to drain out.  The foreline pump and oil may be hot.

        - Replace the drain plug.

        - Remove the book or object used to prop up the pump.

        - Refill the foreline pump with Inland 45 pump oil, using a funnel, until the oil level is between the two fill marks in the site window, approximately 0.28 L of oil.

        - Wait a few minutes for the oil to settle. If the oil level drops, add oil to bring the oil level near the upper line.

        - Reinstall the fill cap.

        - Power on the GC.

        - Ensure that the vent valve is closed.  Holding the MSD chamber door closed, power on the MSD and ensure that the turbo pump speed climbs to 100%.

        - Start the Chemstation software and apply setpoints.

        - Allow the instrument to equilibrate for two hours prior to tuning.

      - Record on the GCMS Maintenance Log.

      - Required post-maintenance check:  Successful daily QCC, refer to 5.6.

    - Clean Source

      - Clean annually, at a minimum.

        - Vent the MSD by selecting the vent option in Instrument Control of Chemstation.  Allow the vent cycle to run, when the cycle is complete and the temperatures are below 100 degrees Celsius turn off the MSD.

        - Press [Oven] and set the oven to 35°C.  Press [Front Inlet] and turn off the inlet temperature and pressure.  When the temperature reaches the setpoint turn the GC off.

        - Open the vent valve

        - Detach the ribbon cables from the circuit board on the MSD chamber door.

        - Pull open the MSD chamber door by hand.

        - Detach the leads from the ion source, loosen the screws and remove the ion source.

        - Remove the filaments using a hex ball driver.

        - Separate the repeller assembly from the source body. The repeller assembly includes the source heater assembly, repeller, and related parts.

        - Remove the repeller.

        - Unscrew the interface socket. A 10-mm open-end wrench fits the flats on the interface socket.

        - Remove the setscrew for the lenses.

        - Push the lenses out of the source body.

        - If insulators are dirty, clean them with a cotton swab dampened with reagent-grade methanol.  If that does not clean the insulators, replace them.  Do not abrasively or ultrasonically clean the insulators.

        - The filaments and source heater assembly cannot be cleaned ultrasonically.  Replace these components if major contamination occurs.

        - Collect the following parts that contact the sample or ion beam to be cleaned.

          - Repeller

          - Interface socket

          - Source body

          - Drawout plate

          - Drawout cylinder

          - Ion focus lens

          - Entrance lens

        - Abrasively clean the surfaces that contact the sample or ion beam.

        - Use an abrasive slurry of alumina powder and methanol on a cotton swab.  Use enough force to remove all discolorations.  Polishing the parts is not necessary; small scratches will not harm performance.  Also abrasively clean the discolorations where electrons from the filaments enter the source body.

        - Rinse away all abrasive residue with reagent-grade methanol.

        - Make sure all abrasive residue is rinsed way before ultrasonic cleaning.  If the methanol becomes cloudy or contains visible particles, rinse again.

        - Separate the parts that were abrasively cleaned from the parts that were not abrasively cleaned.

        - Ultrasonically clean the parts (each group separately) in polar and non-polar solvents; as recommended by instrument manufacturer.

        - Place the parts in a clean beaker.

        - If needed, dry the cleaned parts in an oven at 100 °C for 5–6 minutes.

        - If needed, let the parts cool before you handle them.

        - Take care to avoid recontaminating cleaned and dried parts.  Put on new, clean gloves before handling the parts.  Do not set the cleaned parts on a dirty surface. Set them only on clean, lint-free cloths.

        - Slide the drawout plate and the drawout cylinder into the source body.

        - Assemble the ion focus lens, entrance lens, and lens insulators.

        - Slide the assembled parts into the source body.

        - Install the setscrew that holds the lenses in place.

        - Reinstall the repeller, repeller insulators, washer, and repeller nut into the source heater assembly.  The resulting assembly is called the repeller assembly.

        - Reconnect the repeller assembly to the source body. The repeller assembly includes the source heater assembly, repeller, and related parts.

        - Reinstall the filaments, replace if excessively worn.

        - Reinstall the interface socket.

        - Do not overtighten the repeller nut or the ceramic repeller insulators will break when the source heats up.  The nut should only be finger-tight.

        - Do not overtighten the interface socket.  Overtightening could strip the threads.

        - Reinstall the ion source into the MSD and reattach the leads.

        - Close the vent valve.

        - Push the MSD chamber door close and reattach the ribbon cables to the circuit board.

        - Power on the GC.

        - Holding the MSD chamber door closed, power on the MSD and ensure that the turbo pump speed climbs to 100%.

        - Start the Chemstation software and apply setpoints.

        - Allow the instrument to equilibrate for two hours prior to tuning.

      - Record on the GCMS Maintenance Log.

      - Required post-maintenance check:  Successful daily QCC, refer to 5.6., and monthly QCC, refer to 5.7.

    - Gold Seal

      - Replace annually, at a minimum.

        - Vent the MSD by selecting the vent option in Instrument Control of Chemstation.  Allow the vent cycle to run, when the cycle is complete and the temperatures are below 100 degrees Celsius turn off the MSD.

        - Press [Oven] and set the oven to 35°C.  Press [Front Inlet] and turn off the inlet temperature and pressure.  When the temperature reaches the setpoint turn the GC off.

        - Be careful - The inlet fittings may be hot enough to cause burns.  Loosen the inlet column nut with the ¼ inch wrench and remove the column from the inlet. Cap the open end of the column to prevent contamination.

        - Remove the insulation cup from around the base of the inlet using the star screwdriver.

        - Use the 1/2-inch wrench to loosen the reducing nut, and then remove it.

        - The washer and seal are inside the reducing nut.  Remove them, noting their orientation.

        - Handle the new gold seal and washer with clean, lint-free, non-nylon gloves.  Place the washer in the reducing nut.  Place the new inlet base seal on top of it with the raised portion facing down.

        - Replace the reducing nut. Use the 1/2-inch wrench to tighten the nut.

        - Replace the column and the insulation cup.

        - Using ChemStation load a method to return the GC to appropriate settings.  If prompted, do not save any method changes.

        - Allow the GC to return to the setpoints.

        - Ensure that the vent valve is closed.  Holding the MSD chamber door closed, power on the MSD and ensure that the turbo pump speed climbs to 100%.

        - Start the ChemStation software and apply setpoints.

        - Allow the instrument to equilibrate for two hours prior to tuning.

      - Record on the GCMS Maintenance Log.

      - Required post-maintenance check: Successful daily QCC, refer to 5.6., and monthly QCC, refer to 5.7.

- Procedure

  - Instrument Settings

    - Select a GC-MS method based upon the sample to be analyzed.

      - The 70TOX1 method is used for solid phase extraction fractions.

        - DB5-MS column, 12 m X 0.200 mm X 0.33 µm

        - Initial temperature - 70 oC hold for 1.00 minute

        - 70 oC - 125 oC @ 40 oC/minute hold for 1.00 minute

        - 125 oC - 285 oC @ 17 oC/minute hold for 12.00 minutes

        - Total time of run: 24.79 minutes

        - Scan Range 40 – 500

        - 230 °C source temperature

        - 150 °C quadrupole temperature

      - The 70TOX1Deriv method is used for derivatized basic solid phase extraction fractions.  It has the same parameters as the 70TOX1 method except for an extended scan range to accommodate increased masses.

        - DB5-MS column, 12 m X 0.200 mm X 0.33 µm

        - Initial temperature - 70 oC hold for 1.00 minute

        - 70 oC - 125 oC @ 40 oC/minute hold for 1.00 minute

        - 125 oC - 285 oC @ 17 oC/minute hold for 12.00 minutes

        - Total time of run: 24.79 minutes

        - Scan Range 40 – 550

        - 230 °C source temperature

        - 150 °C quadrupole temperature

      - Each method washes the syringe 10 times between injections to ensure sample integrity

      - When the standard methods are not appropriate to analyze a compound, a modified method may be used in accordance with the CCBI Administrative Procedure for Exceptions.

  - Shutdown / Startup

    - The GC-MS shall be left on at all times.

    - The computer may be shut down or restarted if needed.

    - A successful daily QCC check, refer to 5.6., must be performed following any GC or MS shutdown.

    - Record any shutdown on the GC-MS activity logbook.

  - Place a blank ethyl acetate injection in the sequence immediately before each sample injection as a negative QCC using the same method as the sample.  Additional ethyl acetate blank injections may be placed in the sequence.

    - Evaluate the negative QCC to ensure that the instrument is free of any controlled substance, any substance being identified in the sample and any substance that may interfere with the identification of sample component(s).

      - Record any negative QCC’s that do not meet the requirement in the GC-MS activity log, in the case file and follow the CCBI Laboratory Procedure for Corrective and Preventive Action as required.

      - Note the presence of large amounts of common gas chromatography peaks (e.g., siloxanes) in the GC-MS activity log and notify the DWI Blood Chemistry Technical Leader.

  - Use the current date in the names of sequences.

  - Include the CCBI case file number in the data file name and any additional information needed to uniquely identify the sample.

    - Data files associated with casework and performance checks shall not be deleted or overwritten.

    - The DWI Blood Chemistry Technical Leader shall electronically archive data files annually.  Store archived files electronically on the shared drive.

    - Notify the DWI Blood Chemistry Technical Leader if the data storage location becomes full.

  - The GC-MS provides retention time data and mass spectral data.

  - For full scan methods evaluate the chromatogram and spectra for each peak of interest.

    - Mass Spectral Comparison

      - For a positive mass spectral comparison, the sample mass spectrum must be substantially comparable, i.e., equivalent, to that of primary or secondary reference material.

      - Record in the case file the mass spectrum of the reference material with the supplier and lot number or other DWI Blood Chemistry Unit designation.  Library search results may be included.

    - GC Relative Retention Time (RRT) Comparison

      - For a positive RRT comparison, the percent difference between the sample retention time relative to the internal standard must be within ± 5 % of a primary or secondary reference material relative to the internal standard.  Round relative retention times to the hundredths, refer to 11.7. for rounding instructions.  The retention time of the collected mass spectrum may be used to determine the relative retention time.

      - Record in the case file

        - Sample TIC with the retention time(s) displayed or mass spectrum with retention time displayed.

        - The percent difference of the reference material and sample relative retention times.

    - Record in the case record the:

      - Total Ion Chromatogram (TIC) for the sample and corresponding blank.

        - Mass spectra of significant peaks of interest, including internal standard

        - Expanded mass spectra of any phenethylamines.

      - Batch Negative Control TIC and corresponding blank TIC

        - Batch Negative Control internal standard mass spectrum

        - Extraction ion chromatograms of batch negative control showing absence of any substance identified in sample.

- Calculations

  - RRT Percent Difference Calculation, round to one decimal place, refer to 11.7.:

|[(reference material RT ÷ reference material internal standard RT) – (sample RT ÷ sample internal standard RT)]| ÷ (reference material RT ÷ reference material internal standard RT) * 100

- Limitations

  - The GC-MS methods described in this procedure shall not be used to distinguish between optical isomers.

  - Due to the possibility of conversion of methadone to EDDP (2-ethylidene-1,5-dimethyl-3,3-diphenylpyrrolidine, an inactive metabolite of methadone) during GC/MS analysis do not report EDDP in the presence of methadone.

  - Due to the possibility of conversion of hydromorphone to hydrocodone during GC/MS analysis, do not report hydrocodone in the presence of hydromorphone.

- Safety

  - Refer to the CCBI Health and Safety Manual.

  - Handle syringes with care to avoid punctures.

  - Use extreme caution handling/installing/removing/transporting compressed gas cylinders.  Cylinders shall not be moved without the cylinder cap securely in place.

  - Gas Chromatograph and Mass Spectrometer may be extremely hot.  Avoid touching hot areas and wear protective gloves while performing maintenance.

- References

  - Agilent 7890A GC User Information, Agilent Instrument Utilities V B.1.06.11343.1852.

  - Agilent 5975 MSD User Information, Agilent Instrument Utilities V B.1.06.11343.1852.

  - Agilent 7693A ALS User Information, Agilent Instrument Utilities V B.1.06.11343.1852.

  - Moffat, A. C., et al., eds. Clarke’s Isolation and Identification of Drugs. 2nd Edition. London:  Pharmaceutical Press, 1986.

  - Moffat, A.C., et al., eds. Clarke’s Analysis of Drugs and Poisons. 4rd Edition. London: Pharmaceutical Press, 2011.

  - Pfleger, Maurer, and Weber, Mass Spectral and GC Data of Drugs, Poisons, Pesticides, Pollutants and Their Metabolites; 3rd Ed., Vols. 1-2, 2007.

  - Guide for the Use of the International System of Units (SI). NIST Special Publication 811, 2008 Ed., (March 2008; 2nd printing November 2008). p.43

  - Skoog, Douglas A., F. James Holler and Timothy A. Nieman.  Principles of Instrumental Analysis. 5th Edition. Garcourt Brace & Company, 1998.

  - Agilent GC-MSD ChemStation and Instrument Operation Student Manual Course Number H4043A Volume 1, Revision E.02.xx. Agilent Technologies: printed February 2008.

  - Agilent 5975 Series MSD Operation Manual. 3rd Edition. Agilent Technologies Inc., 2010, Manual Part Number G3170-90036.

  - Agilent G1701EA GC/MSD Chemstation Getting Started. 1st Edition.  Agilent Technologies Inc., 2011, Manual Part Number G1701-90069.

- Records

  - GC-MS Maintenance Log

  - GC-MS Daily QCC Log

  - GC-MS Monthly QCC Log

  - GC-MS Activity Log

  - Case record

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
|  |  | Updated section 3 and line 5.6.5. Added lines 11.10 and 11.11. |
| 1/16/15 | 2 | Additions, minor corrections and clarifications in 4.2.2, 4.2.3, 4.5, 5.3, 5.6.7, 5.7.1, 5.7.4, 5.7.8.2, 6.2, 6.4.3.1, 6.4.5.1, 6.4.6.1, 7.1.1.4, 7.3 and 9.3 |
| 5/12/17 | 3 | Added 5.6.3, 5.6.4, & 5.6.5, Removed portions of 5.6.7.3., 5.6.7.4, Removed section 5.6.7.5 |
| 12/14/2017 | 4 | Updated 4.3.4 to add procedure reference. Updated 5.1, 5.1.1.2, 5.1.3.1, 5.1.3.2, 5.1.3.3, 5.6.2, 5.6.9, 5.7.8, 5.8.4, 6.3.1, 7.5, 7.5.2, to allow for electronic storage. Updated 6.2 to reflect current information recorded on the respective form. Updated 6.4.2.2, 6.4.3.2, 6.4.4.2, 6.4.5.2, 6.4.6.2, 7.2.4, 12.1 to reflect the current form title. Removed 5.1.4 due to electronic format of instrument logbook. Updated 5.2 to include all maintenance. Changed 5.6.1.1 to modify tune requirements. Reorganized 5.6.7 and 5.6.8 to reflect order of performance. Updated 5.7.5,5.7.6, and 5.7.8 to include oxazepam. Updated 6.1 to require the activity log to indicate the instrument is out of service when performing column or mass spectrometer maintenance. Updated 6.3 and 6.3.1 Updated 6.4.5.1.21, 6.4.5.1.23, and 6.4.5.1.24. Updated 6.4.5.1.22. Removed 6.4.1.2. Updated 7.3 to remove requirement for BSTFA blanks. Removed 12.1, 12.2, 12.3, 12.4. Updated 12.5. |
| 4/10/2018 | 5 | Corrected 5.1.3.1 Updated safety manual name in 10.1 |
| 10/28/19 | 6 | Removed CANSIM procedure references in 7.1.1.3, 7.8 and 9.2. |
| 3/23/20 | 7 | 3. Abbreviations - removed. 4.3 Materials - removed DB5 30m column 5.7 Monthly Quality Control Check - Removed reference to “AUTO” Method. 6.4 Suggested Routine Maintenance - added 6.4.2 Merlin microseal 7. Procedure - Removed reference to “AUTO” Method. 9. Limitations - added 9.3 to address hydrocodone and hydromorphone confirmation. |
| 6/10/20 | 8 | 7.7.2. GC Relative Retention Time (RRT) Comparison - removed 7.7.2.2.1 |

# 13: Technical Procedure for DWI Blood Chemistry Analysis

- Purpose/Scope - This procedure provides direction for analysis of blood in DWI submissions in the DWI Blood Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification.

- Definitions

  - Alcohol - Any substance containing any form of alcohol, including ethanol, methanol, propanol, and isopropanol.

  - Alcohol Concentration - The concentration of alcohol expressed as grams of alcohol per 100 milliliters of blood.

The results of a defendant's alcohol concentration determined by a chemical analysis of the defendant's breath or blood shall be reported to the hundredths. Any result between hundredths shall be reported to the next lower hundredth.

  - Impairing Substance - Alcohol, controlled substance under Chapter 90 of the General Statutes, any other drug or psychoactive substance capable of impairing a person's physical or mental faculties, or any combination of these substances.

- Abbreviations

  - approx. - approximately

  - BAC - Blood alcohol concentration

  - conc - concentration

  - c or c/ - containing

  - d - dated

  - diff - difference

  - DWI - Driving while impaired

  - ELISA - Enzyme Linked Immunosorbent Assay

  - ext - extract / extraction

  - GCMS, GC-MS, or GC/MS - gas chromatography–mass spectrometry

  - hs - heat sealed

  - i - initialed

  - init - initial

  - LC - Liquid chromatography

  - LOD - Limit of detection

  - LOQ - Limit of quantitation

  - LTC - labeled to contain

  - me - manila envelope

  - MS - mass spectrum

  - MSD - Mass Selective Detector

  - MRM - Multiple Reaction Monitoring

  - MS-MS - Tandem mass spectrometer

  - NIST - National Institute of Standards and Technology

  - NQCC - negative quality control check

  - NCS - no controlled substances

  - p - page

  - pg - page

  - pb - plastic bag

  - pos - positive

  - PQCC - positive quality control check

  - prelim(s) - preliminary testing(s)

  - QCC - quality control check

  - rgt - reagent

  - rec’d - received

  - ref - reference

  - RRT - relative retention time

  - RT - retention time

  - ret’d - returned

  - sch - schedule

  - std(s) - standard(s)

  - s - sealed

  - TIC - Total ion chromatogram

  - temp - temperature

  - tran(s) - transfer(s)

  - u or us - unsealed

  - w/ or  w - with

  - wt - weight

  - zpb - Ziploc plastic bag

- Equipment, Materials and Reagents

  - Refer to the DWI Blood Chemistry Unit Technical Procedures

- Procedure

  - Store DWI submissions of blood evidence in a refrigerator at 40C (± 30C).

  - Blood may be destroyed as biohazardous waste pursuant to Laboratory Administrative Procedures Manual Chapter 7 and upon notification from the Wake County District Attorney’s Office that the legal conditions allowing destruction of the blood evidence have been satisfied.

  - Have only one item of evidence open for analysis at a time.

  - Maintain materials used in analysis that come in direct contact with evidence in closed containers, cabinets or drawers to avoid contamination.

  - Use disposable glassware, caps and pipette tips for direct contact with evidence, when possible.  Properly dispose of such items immediately after use.

  - All individual evidence analysis vessels must be labeled with identifying information to include the case number.  For plate analysis, a map or diagram shall be created with the identifying information to include the case number.

  - Note the following information for the evidence received in the casefile when available and practical:

    - A description of all packaging, submitting agency item number, collection kit lot number and expiration date, the subject’s name and location on the evidence, the color of the vacutainer stopper(s), and the condition of all seals.

    - A description of the type of fluid within the collection vessel, and any unusual condition of the specimen or collection vessel to include the volume if less than approximately two milliliters of blood, leaks, clots, or microbial growth.

    - Collection vessels used for analysis will be clearly marked with an “A” on the vessel.

  - Retain only the collection vessels if the evidence is to be analyzed. Outer packing may be discarded.

    - Label any paperwork, other than packaging, submitted with the blood evidence with the CCBI case number and initials.  Store the paperwork in the case file as administrative documentation.

    - The collection vessel stopper of the blood tube shall serve as the evidence seal without any additional markings or coverings once the blood tube is received by the Blood Drug Chemist and maintained in their custody.

  - Evaluation of the Examination Request

    - All DWI Blood Chemistry examination requests including request which are unclear or non-specific will be analyzed for blood alcohol concentration (BAC) according to the DWI Blood Chemistry Unit Technical Procedure for Determination of Alcohol, Acetone, and Other Volatiles in Blood by Headspace Gas Chromatography.

    - The DWI Blood Chemistry Unit will perform a drug examination when requested; however, the drug analysis will not be required to be completed when the alcohol concentration is determined to be:

      - Greater than 0.085 gram per 100 milliliter of blood and,

      - The submission form does not indicate a death or serious injury to a victim and,

      - Additional analysis is not requested by the Wake County District Attorney’s Office

      - All drug examinations will consist of examinations for Cannabinoids and basic drugs (and metabolites) using the current DWI Blood Chemistry Unit Technical procedures.

        - An acid/neutral extraction shall be performed for either of the following reasons:

          - If acid/neutral substance(s) are suspected according to the submission form. (Carisoprodol, Barbiturates, etc.)

          - If no controlled substances, controlled substance metabolites or other impairing substance are identified during basic drug and cannabinoid testing and the testing performed has not tested for acid/neutral substances.

    - The screening technique(s) chosen for analysis must include all analytes currently targeted by the laboratory procedures.

      - DWI Blood Chemistry Unit screening techniques include:

        - ELISA Screening

        - LCMS-MS Screening

    - The confirmation technique(s) chosen for analysis must target those drugs or class of drugs indicated in the screen tests.

      - DWI Blood Chemistry Unit confirmation techniques include:

        - LCMS-MS qualitative and quantitative analysis

        - GCMS qualitative analysis

    - The chemist conducting the analysis will determine the extent of the analysis to be completed when there is insufficient blood to complete all requested examinations.  The determination will be based on the information available at the time of the examination.  The reason for not completing an examination as requested will be documented in the case file.  If there is any question as to the extent of the analysis to be completed, the chemist shall contact the requester or the Wake County District Attorney’s Office for further instruction.

  - Identification and Reporting of controlled substance and controlled substance metabolites are required when either of the following analysis conditions are met:

    - Positive screening technique result and a positive confirmation technique result or

    - Two positive confirmation techniques on two separate aliquots

  - Minimum Requirements for Reporting the Absence of Substances

    - Negative Screening technique results and negative confirmation results via GCMS or LC-MS.

  - Reporting

    - The DWI Blood Chemistry Unit Affidavit and Revocation Report form will be completed for all DWI Blood Chemistry Examinations.

    - The report shall include a “Description of Evidence Submitted” field, the submitting agency item number(s), a detailed description of the item(s) and seals and, if applicable, the subject name as it appears on the evidence and a description of where it appears on the evidence in this field.

    - The report shall include a statement of the disposition of the submitted items:

The blood evidence will be retained for a minimum of twelve months from the date this report was notarized.  Following that period the blood evidence may be destroyed without further notification according to North Carolina General Statute §20-139.1.(h).

    - The report shall include a “Results and Conclusions” field.  The results and conclusions of the analysis shall be included in this field.

    - Results and Conclusions Reporting

      - Reporting of Blood alcohol analysis:

        - When only ethanol is to be reported, report the alcohol concentration and uncertainty with the statement:

The blood alcohol concentration is (insert blood alcohol concentration) gram per 100 milliliters of blood, which has been determined as required by NCGS 20-4.01 (1b).

        - When multiple alcohols and/or acetone are to be reported, report the concentration and uncertainty with the statement:

The blood alcohol concentration is (insert blood alcohol concentration) gram per 100 milliliters of blood, which has been determined as required by NCGS 20-4.01 (1b).

Add for each identified alcohol or acetone:

The (insert identified alcohol) concentration is (insert individual alcohol concentration) gram per 100 milliliters of blood, which has been determined as required by NCGS 20-4.01 (1b).

        - Add an additional statement to the report containing the values from the DWI Blood Chemistry Unit Technical Procedure for Uncertainty of Measurement line 4.12

(The average of the four measured alcohol concentrations is (insert blood alcohol concentration) gram per 100 milliliters ± (insert uncertainty) gram per 100 milliliters of blood at a coverage probability of 99.73%.)

        - If no alcohol is identified or if the alcohol concentration is less than 0.01 grams of alcohol per 100 milliliters of blood, report with the statement:

The blood alcohol concentration is 0.00 grams of alcohol per 100 milliliters of blood which has been determined as required by NCGS 20-4.01 (1b).

      - Reporting of analysis for controlled substances and controlled substance metabolites

        - If a controlled substance or controlled substance metabolite is to be reported, report with the statement:

Analysis for the presence of controlled substances and controlled substance metabolites confirmed the presence of:

If a quantitative analysis is performed, include the following statement to the report results:

(The percent uncertainty of measurement(s) listed for quantitative values have a coverage probability of (insert coverage probability).)

(insert the identity of the substance(s), (insert the measured concentration rounded to two significant figures with the appropriate units along with the uncertainty of measurement expressed as a percent))

If quantitative analysis is performed, and the measured quantitative value exceeds the calibration limit, include the following statement in lieu of the measured concentration:

Greater than (insert Calibration limit value) (insert units)

If multiple quantitative results are obtained, the average value shall be reported.

Example:  Analysis for the presence of controlled substances and controlled substance metabolites confirmed the presence of: (The percent uncertainty of measurement(s) listed for quantitative measurements have a coverage probability of 99.73 %)

Delta-9-Tetrahydrocannabinol (THC), 7.2 ng/ml ± 20%

11-Nor-9-Carboxy-Delta 9-THC (Carboxy-THC), a metabolite of THC, 0.25 g/ml ± 10%

Methadone, greater than 0.25 g/ml

5.12.5.2.2.1	If a controlled substance or controlled substance metabolite is identified in one solid phase extract, but there is insufficient quantity of sample to perform a second extraction then the substance may be reported with the statement:

Analysis for the presence of controlled substances and controlled substance metabolites indicated the presence of: (insert the identity of the substance(s)).  Insufficient sample volume was present to perform additional testing to confirm the presence of the substance.

        - If no controlled substances or controlled substance metabolites are to be reported, report with the statement:

Analysis for the presence of controlled substances and controlled substance metabolites did not confirm the presence of any controlled substances or controlled substance metabolites.

        - If a positive indication was observed in the ELISA analysis and the presence of a corresponding controlled substance / metabolite was not confirmed, report with the statement:

The immunoassay screening test for (insert name of immunoassay screening test(s)) gave a positive indication for the presence of (insert name of immunoassay screening test(s)); however, confirmatory testing did not identify drugs indicated by the immunoassay screening test(s).

        - If analysis for a specific controlled substance is requested and the substance cannot typically be identified by the DWI Blood Chemistry Unit Technical Procedures, add to the report:

Analysis for (insert substance(s) requested) was requested; however, (insert substance(s) requested) generally cannot be detected by the current CCBI Crime Laboratory procedures.

        - If in the course of analysis for controlled substances / metabolites a non-controlled substance or metabolite was indicated by GC/MS analysis but the criteria for identification and reporting were not met, add the following statement to the report:

Analysis indicated the presence of a non-controlled substance.

        - If in the course of analysis for controlled substances / metabolites, the identification and reporting requirements for a non-controlled substance / metabolite is met, add to the report:

The presence of the following non-controlled substance(s) / metabolite(s) was confirmed: (insert the identity of the substance(s))

      - Reporting of analysis for non-controlled substances / metabolites

        - If a non-controlled substance / metabolite is to be reported, report with the statement:

Analysis for the presence of non-controlled substances / metabolites confirmed the presence of: (insert the identity of the substance(s))

        - If no non-controlled substances / metabolites are to be reported, report with the statement:

Analysis for the presence of non-controlled substances / metabolites did not confirm the presence of any non-controlled substances / metabolites.

        - If analysis for a specific non-controlled substance is requested and the substance cannot typically be identified by the DWI Blood Chemistry Unit Technical Procedures, add to the report:

Analysis for (insert substance(s) requested) was requested; however, (insert substance(s) requested) generally cannot be detected by the current CCBI Crime Laboratory procedures.

      - Testing Methods

        - Following the result statement(s), include “Method(s):” (insert the methods used for analysis)

          - DWI Blood Chemistry methods for analysis include:

Headspace Gas Chromatography (HSGC), Enzyme Linked Immunosorbent Assay (ELISA), Gas Chromatography Mass Spectrometry (GCMS), Liquid Chromatography Tandem Mass Spectrometry (LCMSMS)

  - Review

    - Upon successful completion of the Administrative and Technical Reviews, sign the Report and have it notarized and disseminated.

- Limitations

  - Refer to the DWI Blood Chemistry Unit technical procedures.

- Safety

  - Refer to the CCBI Health and Safety Manual.

- References

  - Levine, Barry ed., Principles of Forensic Toxicology. 3rd edition. AACC Press, 2009.

  - Moffat, Anthony C. ed. Clarke’s Analysis of Drugs and Poisions, Volume 1, 4th edition, Pharmaceutical Press, 2011.

  - Randall C. Baselt.  Disposition of Toxic Drugs and Chemicals in Man. 8th Ed. (2008).

  - James C. Garriott (Editor), Medicolegal Aspects of Alcohol, 5th  Ed., 2008.

  - Baselt, Randall C. Drug Effects on Psychomotor Performance, Foster City, California: Biomedical Publications, 2001.

  - North Carolina General Statutes.

https://www.ncleg.gov/Laws/GeneralStatutes

  - Forensic Toxicology Laboratory Guidelines. SOFT/AAFS. 2006.

- Records

  - CCBI Laboratory Technical and Administrative Review/Coversheet

  - Case Record

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 2/8/13 | 1 | Compliance with ASCLD/LAB requirements |
| 11/8/13 | 2 | Incorporation of DBCTP14 |
| 7/14/14 | 3 | Include use of Agency item numbers and ref to LAPM19 |
| 1/16/15 | 4 | Updates, additions, clarifications and minor corrections in 2.3, 5.2, 5.2.1, 5.2.1.1, 5.2.1.2, 5.2.2, 5.2.2.2, 5.2.3, section 5.2.4, section 5.4.1, 5.7.4, 5.7.5, 5.9.3.7, 5.12.3, 5.12.5.1.1, 5.12.5.1.2, 5.12.5.1.3, 5.12.5.1.4, 5.12.5.2.6, 5.12.5.2.7, 9.2, and clarification of Blood Drug Chemist throughout. |
| 8/31/15 | 5 | Update wording for non-controlled substances. Remove wording regarding District Attorney’s Office requests for additional analysis from laboratory reports. |
| 11/1/16 | 6 | Added wording to allow reporting substances identified in one sampling as indicated in 5.9.3.5, 5.10.2, and 5.12.5.2.2.1. Corrected spelling mistake in Revision History Version 5. Added statement in 5.9.3.4 to consider controlled substance metabolites as indicated by ELISA. |
| 9/15/17 | 7 | Removed requirement in 5.2.1.2 Added an option for other evidence location in 5.2.2.2 Updated 5.2.4.3 Removed requirement for weekly cleaning form in 5.4.1.7 and 9.2 Added amphetamine to the ELISA screen panel in 5.9.3.2 Removed the uncertainty of measurement statement for the statutorily required reporting value in 5.12.5.1.1 and 5.12.5.1.2 and added NCGC reference to reporting statement. Updated 5.12.5.1.3, removed qualifier statement in 5.12.5.1.4 Removed requirement for left hand corner in 5.12.6.1.2.1 and 5.12.6.2.2.1 Corrected references in 5.13.3.3.2, 5.13.3.3.3, 5.13.4.3, and 5.13.4.3.1 |
| 1/26/18 | 8 | Updated 5.13.2 and 5.13.3. to conform with changes in Laboratory Administrative Procedures Manual, removed 5.13.4. to conform with changes in Laboratory Administrative Procedures Manual. |
| 4/10/2018 | 9 | Updated safety manual name in 7. |
| 7/8/19 | 10 | Removed 5.6 due to inclusion into LAPM |
| 10/28/19 | 11 | Updated Cannabinoid reporting procedures in 5.8.3.7, 5.9.3, 5.11.5.2.1, 5.11.5.2.2, 5.11.5.3.3, 5.11.5.2.4. |
| 01/23/20 | 12 | 3. Abbreviations - added LC, LOD, LOQ, MSD, MRM, MS-MS, and TIC. 5. Procedure - revised section and removed any instructions already in other CCBI administrative procedures. Added 5.7.3. requirement to label the sample vessel used for analysis with “A”. Removed additional reports in 5.12.6. 8. References - added 8.7. |
| 7/08/20 | 13 | 5.9. Evaluation of the Examination Request - revised sections to be clearer and added scenario when an acid extraction will be performed. 5.9.5 - Removed section addressing examples of positive results 5.10. Minimum Requirements for Identification and Reporting of controlled substance and controlled substance metabolites - Specified the elevated ELISA assays for identification in 5.10.1. 5.11. Minimum Requirements for Reporting the Absence of Substances - added negative LC-MS analysis 5.12.5.2. Reporting of analysis for controlled substances and controlled substance metabolites - added report wording to be used when analysis results exceed the calibration limit and when multiple quantitative results are obtained 5.12.5.2.1. Added criteria for reporting results from multiple quantitative results |
| 9/8/20 | 14 | 5.9.2.4.1.2. - added clarifying statement regarding acid/neutral analysis. 5.12.5.4. Testing Methods - added section to include testing method on reports |
| 10/26/20 | 15 | 5.10 Requirements for reporting controlled substances and controlled substance metabolites - revised wording |
| 01/20/21 | 16 | 5.10 Requirements for reporting controlled substances and controlled substance metabolites - revised wording |

# 14: Technical Procedure for Solid Phase Extraction of Benzodiazepines for GC-MS Analysis

- Purpose / Scope - The procedure is used to isolate benzodiazepine drugs from blood using solid phase extraction for analysis by gas chromatography - mass spectrometry in the DWI Blood Chemistry Unit of the CCBI Crime Laboratory.

- Definitions

  - Quality control check – Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

  - QCC – Quality control check

- Equipment, Materials and Reagents

  - Equipment

    - pH meter

    - TurboVap Evaporator

    - Mechanical pipettors and corresponding tips

    - Volumetric flasks, class A

    - Volumetric pipets, class A

  - Materials

    - Test tubes, 16 x 125mm and 13 x 100mm

    - Test tube caps or stoppers

    - Vortex Mixer

    - Centrifuge

    - UCT Clean Screen DAU Solid Phase Extraction Columns

    - Deionized water

    - Positive pressure manifold

    - Nitrogen, UHP

  - Reagents

    - Critical Reagents

      - Negative blood

      - BSTFA with 1% TMCS (N,O-bis(trimethylsilyl)trifluoroacetamide with 1% trimethylchlorosilane)

    - Reference materials

      - Prazepam

    - Commercial reagents

      - Ammonium Hydroxide, ACS grade,

      - Ethyl Acetate, ACS grade

      - Methanol, ACS grade

      - Acetonitrile, ACS grade

      - Disodium Hydrogen Phosphate, Anhydrous, ACS grade

      - Sodium Dihydrogen Phosphate, Monohydrate, ACS grade

    - Prepared Reagents

      - 20 % Acetonitrile in 100mM Phosphate Buffer

        - Thoroughly mix acetonitrile with 100 mM phosphate buffer to yield the desired volume of a solution containing 20 % acetonitrile.  Example:  5 ml acetonitrile and 95 ml 100 mM phosphate buffer.

        - Lot number:    Eight digit format year/month/day/20%CH3CN:95%Buf/initials of preparer.

        - Expiration:  end of day

        - Storage:  closed container

        - QCC:  successful negative control extraction

      - Ethyl Acetate with 4% Ammonium Hydroxide

        - Thoroughly mix ethyl acetate with ammonium hydroxide to yield the desired volume of solution containing 4% ammonium hydroxide.  Example:  96 ml ethyl acetate and 4 ml ammonium hydroxide.

        - Lot number:    Eight digit format year/month/day/AmmEtOAc/initials of preparer.

        - Expiration:  end of day

        - Storage:  closed container

        - QCC:  basic to litmus or pH paper

      - 0.1 M Monobasic sodium phosphate, NaH2PO4

        - Dissolve 1.38 grams monobasic sodium phosphate, monohydrate, in  deionized water in a 100 mL volumetric flask.

        - Mix and dilute to 100 mL with deionized water.

        - Lot number:  Eight digit format year/month/day/0.1M_NaH2PO4/initials of preparer

        - Expiration:  one year

        - Storage:  closed container

        - QCC:    Tests acidic with pH or litmus paper

      - 0.1 M Dibasic sodium phosphate, Na2HPO4

        - Dissolve 1.42 grams dibasic sodium phosphate, anhydrous, in 80 mL deionized water in a 100 mL volumetric flask.

        - Mix and dilute to 100 mL with deionized water.

        - Lot number:  Eight digit format year/month/day/0.1M_Na2HPO4/initials of preparer.

        - Expiration:  one year

        - Storage:  closed container

        - QCC:  Tests basic with pH paper or litmus paper.

      - 100mM Phosphate Buffer, pH 6.0

        - Dissolve 0.85 g Na2HPO4 and 6.07 g NaH2PO4-H2O in 400 mL DI water and dilute to 500mL.

        - Adjust pH to 6.0 +/- 0.1 with 0.1M monobasic sodium phosphate (lowers pH) or 0.1M dibasic sodium phosphate (raises pH).

        - Lot number:  Eight digit format year/month/day/0.1MPhosBuffer/initials of preparer

        - Storage:  refrigerate in closed container

        - Expiration:  one month

        - QCC:  Record final pH

      - Prazepam Internal Standard Stock Solution, 100 μg/ml

        - Dilute 1.0 ml of a 1.0 mg/ml prazepam reference standard solution to 10 ml in a class A volumetric flask.  Mix.

        - Lot number:  Eight digit format year/month/day/100μg/mlPrazepamISStock/initials of preparer.

        - Storage:  freezer in closed container

        - Expiration:  three years

        - QCC:  N/A, refer to 4.3.4.7

      - Prazepam Internal Standard Solution, 2μg/ml

        - Pipet 1.0 ml Prazepam Internal Standard Stock solution, 100 ng/ml, into a 50 ml class A volumetric flask.  Dilute to volume with methanol.

        - Lot number:  Eight digit format year/month/day/2μg/mlPrazepamIS /initials of preparer

        - Storage:  refrigerate in closed container

        - Expiration:  one year

        - QCC:  successful negative control extraction

- Quality Control

  - Positive control

    - Prazepam is added to each basic extraction sample as an internal standard.  For each basic extraction sample, the mass spectrum of the prazepam internal standard must have a positive comparison to prazepam reference material and the signal-to-noise ratio for the prazepam internal standard gas chromatographic peak must be 5:1 or greater.  The signal-to-noise ratio is defined as the response at the peak height of the internal standard ÷ the response at the baseline or valley immediately before the internal standard signal.

  - Negative Control

    - For each extraction batch of blood samples prepare a negative control as directed in Section 6 with 2.0 mL of negative blood.

    - The mass spectrum of the prazepam internal standard of the negative control must have a positive comparison to prazepam reference material and the signal-to-noise ratio for the prazepam internal standard gas chromatographic peak must be 5:1 or greater.  The signal-to-noise ratio is defined as the response at the peak height of the internal standard ÷ the response at the baseline or valley immediately before the internal standard signal.

    - The negative control must not contain any controlled substances, controlled substance metabolites or any substance being identified in the sample.  The negative control shall be subjected to the same post extraction techniques as any case samples in the batch.

- Procedure

  - Allow all solutions and samples to equilibrate to room temperature.

  - Ensure that all blood samples are homogenous by shaking and/or vortexing.

    - If a homogenous sample cannot be obtained, make a notation in the worksheet detailing the condition of the sample and its handling.

  - To 1 mL of 100 mM phosphate buffer add 100 μl of the prazepam internal standard solution, mix/vortex.

  - Add 2.0 mL of blood sample to be analyzed (case sample or negative blood for the negative control) using a mechanical pipettor labeled “blood” and add 3 mL of deionized water.

  - Mix/vortex and let stand for 5 minutes to lyse red blood cells.

  - Mix/vortex sample and centrifuge for 10 minutes at >2000 RPM

  - Decant liquid portion of the sample into 2 mL of 100 mM phosphate buffer solution.

  - Place the Clean Screen DAU Extraction columns into the positive pressure manifold column rack and place the column rack onto the support arms of the manifold.

  - Turn on the nitrogen to the manifold and select the corresponding switches on the front and rear of the manifold to allow nitrogen flow to the corresponding columns.

  - Adjust the two pressure regulators on the front of the manifold.  Use the low pressure regulator to load or aspirate a sample or liquid from the column at a flow rate of 1 to 2 ml per minute, 1 psi or less.  Use the high pressure regulator to dry the column with a pressure setting of 80 psi.

  - Condition the column with 3 ml methanol.

  - Rinse the column with 3 ml of Ethyl Acetate with 4% Ammonium Hydroxide

  - Condition the column with 3 ml deionized water

  - Condition the column with 1 ml 100 mM phosphate buffer

  - Load sample onto column

  - Rinse column with 3 ml deionized water

  - Rinse column with 3 ml of 20 % Acetonitrile in 100mM Phosphate Buffer

  - Dry column for 5 minutes at 80 psi nitrogen

  - Rinse column with 2 ml hexane

  - Collect benzodiazepines with 3 ml of Ethyl Acetate with 4% Ammonium Hydroxide

  - Evaporate the solvent from the collection test tube using the TurboVap LV Evaporator, refer to the DWI Blood Chemistry Unit Technical Procedure for Biotage TurboVap LV Evaporator.  Remove the tubes immediately upon reaching dryness.

  - Reconstitute with 50 µl ethyl acetate or perform silylation as directed in 6.22.1 through 6.22.4 for improved analysis by GC-MS.

    - Derivatize the extract in the collection tube or transfer to a vial using ethyl acetate and evaporate the ethyl acetate. Discontinue evaporation immediately upon reaching dryness

    - Derivatize by adding 50 μl of BSTFA with 1% TMCS to the vial or test tube and capping.

    - Mix and heat the vial or test tube for 30 minutes at 80°C.

    - Remove from the heat source and allow the vial or test tube to cool.

  - Analyze by GC-MS, refer to the DWI Blood Chemistry Unit Technical Procedure for Gas Chromatography – Mass Spectrometry.

  - Record the following in the case file

    - Lot number of internal standard

    - Lot number of negative blood

    - Lot number of UCT Clean Screen DAU Solid Phase Extraction Columns

    - If applicable, the lot number of BSTFA with 1 % TMCS

- Limitations

  - Refer to the references and other published chemical information as needed to determine the need for derivatization.  Some benzodiazepines may need to be derivatized for detection by GC-MS, e.g., a benzodiazepine ELISA indication and no corresponding substance detected in a non-derivatized sample.

  - Do not allow the solid phase extraction columns dry during the extraction other than at steps indicated.

  - Store solid phase extraction columns in a closed container.

- Safety

  - Refer to the CCBI Health and Safety Manual

- References

  - Benzodiazepines in Serum or Plasma for HPLC Analysis Using 200 mg Clean Screen Extraction Column, United Chemical Technologies Inc. Bristol, PA., (2008) 9 –12, 56 – 58.

  - BSTFA with 1 % TMCS Product Specification, Sigma-Aldrich Co, (1997).

  - Moffat, Anthony C. ed. Clarke’s Analysis of Drugs and Poisions, Volume 1, 4th edition, Pharmaceutical Press, 2011.

  - O’Neal, Maryadele J. ed. The Merck Index – An Encyclopedia of Chemicals, Drugs and Biologicals, Merck & Co Inc., Whitehouse Station, NJ, (2006).

  - Benzodiazepines in Serum or Plasma for HPLC Analysis Using:  200 mg Clean Screen Extraction Column, Internal Publication, United Chemical Technologies Inc. Bristol, PA., (2008).

  - Benzodiazepines in Blood Confirmations Using:  200 mg Clean Screen Extraction Column, Internal Publication, United Chemical Technologies Inc. Bristol, PA., (2008).

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 11/8/13 | 1 | Compliance with ASCLD/LAB requirements |
| 1/16/15 | 2 | Updates and corrections to 4.3.4.1.1, 4.3.4.3.1, 4.3.4.6.5, 4.3.4.7.1, 5.1.1, 5.2.2, 6.17 and 6.22.1. |
| 4/10/2018 | 3 | Updated safety manual name in 8. |

# 15: Technical Procedure for Liquid Chromatography/Mass Spectrometry (LCMS-MS)

- Purpose / Scope - This procedure provides direction for the initial setup, performance checks and usage of liquid chromatograph – tandem mass spectrometer instruments in the DWI Blood Chemistry Unit of the Raleigh/Wake City-County Bureau of Identification Crime Laboratory.

- Definitions

Molecular Ion - An ion produced when a molecule gains or loses an electron.  Also referred to as a parent or precursor ion.

Multiple Reaction Monitoring (MRM) - A triple quadrupole instrument method in which a precursor ion is passed through the first quadrupole into the second quadrupole which is a collision cell filled with a collision gas.  The precursor ion is fragmented in the collision and a specific product ion is passed through the third quadrupole for analysis.

Performance verification - The initial confirmation of the reliability of a previously or externally validated method or instrument.

Product ion - An ion produced from the collision of the precursor (parent) ion with a collision gas in the collision cell. Also referred to as a daughter ion.

Quality control check (QCC) - Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

Reference Material - Material sufficiently homogeneous and stable with reference to specified properties, which has been established to be fit for its intended use in measurement or in examination of nominal properties.

Tuning - The process of achieving optimal mass scale and peak intensities by optimizing voltages under a specific set of operating conditions.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

- Equipment, Materials and Reagents

  - Equipment

    - Shimadzu liquid chromatograph with Shimadzu 8050 Tandem Mass Spectrometer with automatic liquid sampler and tray

    - Nitrogen Generator

    - Computer with Shimadzu LabSolutions software and printer

  - Reference Materials

    - Polyethylene glycol (PEG) tuning standard

  - Materials

    - Phenomenex Kinetex Phenyl-Hexyl 4.6 x50mm, 2.6µm and the associated guard column

    - Thermo Scientific Accucore Biphenyl 50 x2.1mm and the associated guard column

    - UCT Selectra DA 100 x 2.1mm, 3um and the associated guard column

    - Autosampler needles

    - Desolvation lines

  - Commercial Reagents

    - Methanol, LCMS grade

    - Isoproponal, LCMS grade

    - Argon gas, Grade 5.0

  - Prepared Reagents

    - Autosampler Needle Wash Solution

      - Mix 750 ml of Methanol with 250 ml of deionized water and place in the Autosampler needle wash solution bottle.

      - Lot Number:  Eight digit format year/month/day/NeedleWash/initials of preparer.

      - Expiration:  one month

      - Storage:   closed container

      - QCC:  None

- Standards and Controls

  - An LCMS-MS logbook shall be maintained.  The logbook shall contain the LCMS-MS Activity Log, the LCMS-MS Maintenance Log and any manufacturer’s certificates, performance verification documentation, QCC printouts and maintenance documentation.

    - LCMS-MS Activity Log

      - Record the date, sequence identification, initials of operator, column, method, and any comments for each sequence analyzed on the DWI Blood Chemistry LCMS-MS Activity Log.

      - Upon completion of each sequence, maintain the sequence in the logbook.

      - Record any error messages on the DWI Blood Chemistry LCMS-MS Activity Log.

  - The Blood Drug Chemist shall record any malfunctions or error messages in the LCMS-MS Activity Log, notify the DWI Blood Chemistry Technical Leader of any malfunctions or error messages and place the instrument out of service by marking the LCMS-MS Activity Log “Out of Service.”

  - The DWI Blood Chemistry Technical Leader shall correct any problems with the instrument or request service.  The DWI Blood Chemistry Technical Leader shall examine the effect(s), if any, of a malfunction or error message on analysis results and implement the CCBI Laboratory Procedure for Corrective and Preventive Action as required.

  - Tuning - A tuning calibration shall be done every 6 months or additionally as needed using the PEG tuning solution.

    - Perform the tune as described in the Shimadzu LCMS 8050 Instruction Manual.

    - Record the tune in the instrument log.  If a problem is noted, notify the DWI Blood Chemistry Technical Leader to determine the cause and appropriate course of action.

  - Performance Verification for New Instrumentation

    - New LCMS-MS instruments shall be installed by a manufacturer representative or approved vendor according to the manufacturer’s guidelines.

- Maintenance

  - Place the instrument out of service prior to performing any column or mass spectrometer maintenance (other than a column change or a cleaning of the cone), by marking the LCMS-MS activity log “Out of Service.”  When the instrument is ready to be returned to service, mark the LCMS-MS activity log “Back in Service.”

  - Record all maintenance on the LCMS-MS maintenance log at the time it is performed with the name of the person performing the maintenance or repairs, the initials of the Blood Drug Chemist recording the maintenance or repairs, the date, the time, a description of the maintenance or repairs and a list of any parts replaced, the type of post maintenance QCC performed, and any additional comments.

    - The DWI Blood Chemistry Technical Leader shall update the maintenance log when the instrument is ready to be used for casework and save any generated data in the instrument logbook folder.

  - Suggested Routine Maintenance - The routine maintenance schedule is a suggested guideline.  The maintenance schedule will be determined by the DWI Blood Chemistry Technical Leader based upon instrument usage and performance.

    - Syringe Needle

      - Replace when a lack of performance is noted.

      - Required post-maintenance check:  Successful QC standard.

    - Desolvation Line

      - Replace as needed.

      - Required post-maintenance check:  Successful QC standard.

    - Clean Source Cone

      - Clean the source cone weekly when in use, at a minimum.

        - Wipe the plate and interior surface with deionized water.

        - Wipe the plate and interior surface with LC/MS Grade Methanol.

        - Required post-maintenance check: Successful system check.

- Procedure

  - Select a LCMSMS screening or quantitation method based upon the type of analysis to be performed.

    - The THC Screen method is used to screen for Cannabinoids.

      - Mobile Phase A is 0.1% Formic Acid in Deionized Water (see DBCTP11 for details of preparation).

      - Mobile Phase B is 0.1% Formic Acid in Methanol (see DBCTP11 for details of preparation).

      - THC Screen LC Gradients

| Time (min.) | % A | % B | Flow Rate (μL/min.) |
| --- | --- | --- | --- |
| 0 | 85 | 15 | 500 |
| 5.2 | 10 | 90 | 500 |
| 6.5 | 10 | 90 | 500 |
| 6.51 | 85 | 15 | 500 |
| 7.5 | 85 | 15 | 500 |

      - THC Screen LCMS-MS Conditions

| Column types | Thermo Scientific Accucore Biphenyl 50x2.1mm |
| --- | --- |
| Mass spectrometer mode | Positive electrospray ionization, dynamic multiple reaction monitoring (dMRM) |
| Injection Volume | 10 µl |
| Column Heater | Ambient ºC |
| Interface Temperature | 350 ºC |
| Heating Gas Flow | 10 L/min [Nitrogen] |
| Nebulizer Gas Glow | 3 L/min [Nitrogen] |
| DL Temperature | 300 ºC |
| Heat Block Temperature | 400 ºC |

      - THC Screen MRM parameters

| Compound | Parent Ion | Product 1 |
| --- | --- | --- |
| THC | 315.46 | 193.00 |
| THC-COOH | 345.46 | 327.00 |
| THC-OH | 331.47 | 313.00 |
| THC-D3 | 318.20 | 122.95 |
| THC-COOH-D9 | 354.20 | 336.25 |
| THC-OH-D3 | 334.20 | 316.10 |

    - The THC Quant method is used for quantitating Cannabinoids.

      - Mobile Phase A is 0.1% Formic Acid in deionized Water (see DBCTP11 for details of preparation).

      - Mobile Phase B is 0.1% Formic Acid in Acetonitrile (see DBCTP11 for details of preparation).

      - THC Quantitation LC Gradients

| Time (min.) | % A | % B | Flow Rate (μL/min.) |
| --- | --- | --- | --- |
| 0 | 45 | 55 | 600 |
| 3.5 | 45 | 55 | 600 |
| 3.6 | 20 | 80 | 600 |
| 5.5 | 20 | 80 | 600 |
| 5.6 | 45 | 55 | 600 |
| 8.0 | 45 | 55 | 600 |

      - THC Quantitation LCMS-MS conditions

| Column types | UCT Selectra DA 100 x 2.1mm, 3um |
| --- | --- |
| Mass spectrometer mode | Positive electrospray ionization, dynamic multiple reaction monitoring (dMRM) |
| Injection Volume | 10 µl |
| Column Heater | 50 ºC |
| Interface Temperature | 350 ºC |
| Heating Gas Flow | 10 L/min [Nitrogen] |
| Nebulizer Gas Glow | 3 L/min [Nitrogen] |
| DL Temperature | 250 ºC |
| Heat Block Temperature | 400 ºC |

      - THC Quantitation MRM parameters

| Compound | Parent Ion | Product 1 | Product 2 |
| --- | --- | --- | --- |
| THC | 315.46 | 193.00 | 123.00 |
| THC-COOH | 345.46 | 327.00 | 299.00 |
| THC-OH | 331.47 | 313.00 | 193.00 |
| THC-D3 | 318.20 | 122.95 | - |
| THC-COOH-D9 | 354.20 | 336.25 | - |
| THC-OH-D3 | 334.20 | 316.10 | - |

    - The MultiDrug Screen method is a screening method for basic drugs.

      - Mobile Phase A is 0.1% Formic Acid in Deionized Water (see DBCTP16 for details of preparation).

      - Mobile Phase B is 0.1% Formic Acid in Methanol (see DBCTP16 for details of preparation).

      - MultiDrug Screen LC Gradients

| Time (min.) | % A | % B | Flow Rate (μL/min.) |
| --- | --- | --- | --- |
| 0 | 85 | 15 | 500 |
| 5.2 | 10 | 90 | 500 |
| 5.6 | 10 | 90 | 500 |
| 5.61 | 85 | 15 | 500 |
| 7.6 | 85 | 15 | 500 |

      - MultiDrug Screen LC Conditions

| Column types | Thermo Scientific Accucore Biphenyl 50x2.1mm |
| --- | --- |
| Mass spectrometer mode | Positive electrospray ionization for basic drugs, Negative electrospray ionization for acidic drugs, dynamic multiple reaction monitoring (dMRM) |
| Injection Volume | 5 µl |
| Column Heater | Ambient |
| Interface Temperature | 350 ºC |
| Heating Gas Flow | 10 L/min [Nitrogen] |
| Nebulizer Gas Glow | 3 L/min [Nitrogen] |
| DL Temperature | 250 ºC |
| Heat Block Temperature | 400 ºC |

      - MultiDrug Screen MRM parameters

| Compound | Parent Ion | Product 1 |
| --- | --- | --- |
| 7-Aminoclonazepam | 286.10 | 121.00 |
| 4-ANPP | 281.10 | 188.05 |
| 6-MAM | 328.00 | 165.00 |
| Alfentanil | 417.20 | 197.00 |
| Alpha-PVP | 232.05 | 90.90 |
| Alprazolam | 309.30 | 281.00 |
| Amphetamine | 136.00 | 90.90 |
| Benzoylecgonine | 290.05 | 168.05 |
| Buprenorphine | 468.30 | 396.00 |
| Butalbital | 223.00 | 42.00 |
| Carfentanil | 395.30 | 335.00 |
| Carisoprodol | 261.05 | 176.10 |
| Chlordiazepoxide | 300.25 | 226.95 |
| Clonazepam | 316.20 | 270.00 |
| Cocaethylene | 318.20 | 195.95 |
| Cocaine | 304.05 | 182.00 |
| Codeine | 300.10 | 215.00 |
| Dextromethorphan | 272.10 | 215.05 |
| Diazepam | 285.25 | 193.05 |
| Dibutylone | 236.10 | 85.95 |
| Dihydrocodeine | 302.10 | 199.00 |
| EDDP | 279.00 | 235.05 |
| Estazolam | 295.10 | 266.95 |
| Etizolam | 343.10 | 313.85 |
| Fentanyl | 337.20 | 188.10 |
| Flunitrazepam | 314.00 | 268.10 |
| Flurazepam | 388.40 | 315.00 |
| Hydrocodone | 300.10 | 199.00 |
| Hydromorphone | 286.05 | 185.00 |
| Ketamine | 238.25 | 124.95 |
| Lorazepam | 322.90 | 276.90 |
| MDA | 180.10 | 162.90 |
| MDEA | 209.00 | 164.00 |
| MDMA | 195.10 | 163.95 |
| MDPV | 277.10 | 206.00 |
| Meperidine | 249.10 | 221.05 |
| Meprobamate | 219.05 | 158.05 |
| Methadone | 310.20 | 265.00 |
| Methamphetamine | 150.05 | 90.95 |
| Methylone | 208.05 | 160.05 |
| Methylphenidate | 234.10 | 83.90 |
| Midazolam | 326.30 | 291.10 |
| Morphine | 286.05 | 152.90 |
| N-Ethylpentalone | 250.10 | 201.95 |
| Nitrazepam | 282.00 | 236.05 |
| N-methyl-norfentanyl | 247.10 | 97.95 |
| Nordiazepam | 271.30 | 139.95 |
| Norhydrocodone | 286.05 | 198.95 |
| Noroxycodone | 301.95 | 283.95 |
| Noroxymorphone | 288.00 | 270.00 |
| O-desmethyl-tramadol | 251.10 | 57.95 |
| Oxazepam | 288.90 | 243.00 |
| Oxycodone | 316.10 | 298.10 |
| Oxymorphone | 302.05 | 283.95 |
| Pentobarbital | 225.00 | 182.05 |
| Phenazepam | 350.90 | 206.00 |
| Phencyclidine | 245.10 | 85.95 |
| Phenobarbital | 231.00 | 42.00 |
| Phentermine | 150.05 | 90.90 |
| Prazepam | 325.40 | 271.1 |
| Secobarbital | 237.00 | 194.00 |
| Tapentadol | 222.15 | 106.95 |
| Temazepam | 301.25 | 255.15 |
| Tramadol | 265.10 | 58.00 |
| Zolpidem | 308.10 | 235.10 |
| 4-ANPP-D5 | 286.25 | 187.95 |
| 6-MAM-D6 | 334.10 | 165.00 |
| 7-Aminoclonazepam-D4 | 290.00 | 121.00 |
| Alpha-PVP-D8 | 241.00 | 90.90 |
| Alprazolam-D5 | 314.05 | 285.90 |
| Amphetamine-D11 | 147.10 | 98.00 |
| Benzoylecgonine-D8 | 298.10 | 170.95 |
| Buprenorphine-D4 | 472.15 | 58.95 |
| Butylone-D3 | 224.90 | 176.95 |
| Carfentanil-D5 | 400.00 | 340.00 |
| Carisoprodol-D7 | 268.15 | 183.10 |
| Chlordiazepoxide-D5 | 305.00 | 288.00 |
| Clonazepam-D4 | 320.00 | 273.95 |
| Cocaethylene-D3 | 321.20 | 198.95 |
| Cocaine-D3 | 307.00 | 184.95 |
| Codeine-D6 | 306.10 | 217.95 |
| Dextromethorphan-D3 | 275.15 | 215.00 |
| Diazepam-D5 | 290.00 | 198.05 |
| Dihydrocodeine-D6 | 308.10 | 202.00 |
| EDDP-D3 | 282.10 | 234.95 |
| Estazolam-D5 | 300.30 | 272.05 |
| Etizolam-D3 | 346.00 | 316.95 |
| Fentanyl-D5 | 342.15 | 104.90 |
| Flunitrazepam-D7 | 321.05 | 274.95 |
| Hydrocodone-D6 | 306.10 | 202.00 |
| Hydromorphone-D6 | 292.10 | 184.90 |
| Ketamine-D4 | 242.00 | 128.95 |
| MDA-D5 | 185.05 | 168.05 |
| MDEA-D6 | 214.10 | 166.05 |
| MDMA-D5 | 200.05 | 166.05 |
| MDPV-D8 | 285.10 | 135.05 |
| Meperidine-D4 | 252.10 | 178.00 |
| Methadone-D9 | 319.20 | 267.95 |
| Methamphetamine-D11 | 161.00 | 96.90 |
| Methylone-D3 | 212.00 | 164.05 |
| Methylphenidate-D9 | 243.00 | 92.95 |
| Midazolam-D4 | 330.00 | 294.95 |
| Morphine-D6 | 292.10 | 152.00 |
| Nordiazepam-D5 | 276.00 | 139.95 |
| Norhydrocodone-D3 | 289.10 | 202.00 |
| Noroxycodone-D3 | 306.10 | 287.95 |
| O-desmethyl-tramadol-D6 | 256.10 | 64.00 |
| Oxazepam-D5 | 292.00 | 245.90 |
| Oxycodone-D6 | 322.10 | 303.95 |
| Oxymorphone-D6 | 305.05 | 287.05 |
| Phenazepam-D4 | 354.90 | 209.95 |
| Phencyclidine-D5 | 249.15 | 96.00 |
| Phentermine-D5 | 155.15 | 96.00 |
| Prazepam-D5 | 330.35 | 275.95 |
| Tapentadol-D3 | 225.10 | 106.95 |
| Temazepam-D5 | 306.00 | 259.95 |
| Tramadol-13C-D3 | 268.15 | 57.90 |
| Zolpidem-D6 | 314.10 | 234.90 |
| Butalbital-D5 | 228.00 | 42.05 |
| Pentobarbital-D5 | 230.00 | 186.95 |
| Phenobarbital-D5 | 235.95 | 42.05 |
| Secobarbital-D5 | 242.00 | 199.00 |

    - The MultiDrug Quant method is a quantitation method for basic drugs.

      - Mobile Phase A is 10 mM Ammonium Formate in deionized Water (see DBCTP16 for details of preparation).

      - Mobile Phase B is 0.1% Formic Acid in Methanol (see DBCTP16 for details of preparation).

      - MultiDrug Quant LC gradients

| Time (min.) | % A | % B | Flow Rate (μL/min.) |
| --- | --- | --- | --- |
| 0 | 95 | 5 | 500 |
| 4 | 0 | 100 | 500 |
| 5 | 0 | 100 | 500 |
| 5.1 | 95 | 5 | 500 |
| 7 | 95 | 5 | 500 |

      - MultiDrug Quant LCMS-MS conditions

| Column types | Phenomenex Kinetex Phenyl-Hexyl 4.6x50mm, 2.6µm |
| --- | --- |
| Mass spectrometer mode | Positive electrospray ionization for basic drugs, Negative electrospray ionization for acidic drugs, dynamic multiple reaction monitoring (dMRM) |
| Injection Volume | 5 µl |
| Column Heater | 35 ºC |
| Interface Temperature | 350 ºC |
| Heating Gas Flow | 10 L/min [Nitrogen] |
| Nebulizer Gas Glow | 3 L/min [Nitrogen] |
| DL Temperature | 300 ºC |
| Heat Block Temperature | 400 ºC |

      - MultiDrug Quant MRM parameters

| Compound | Parent Ion | Product 1 | Product 2 |
| --- | --- | --- | --- |
| 7-Aminoclonazepam | 286.10 | 121.00 | 222.00 |
| 4-ANPP | 281.10 | 188.05 | 105.00 |
| 6-MAM | 328.00 | 165.00 | 211.00 |
| Alfentanil | 417.20 | 197.00 | 268.05 |
| Alpha-PVP | 232.05 | 90.90 | 126.05 |
| Alprazolam | 309.30 | 281.00 | 274.05 |
| Amphetamine | 136.00 | 90.90 | 118.95 |
| Benzoylecgonine | 290.05 | 168.05 | 104.90 |
| Buprenorphine | 468.30 | 396.00 | 414.00 |
| Butalbital | 223.00 | 42.00 | 179.95 |
| Carfentanil | 395.30 | 335.00 | 113.00 |
| Carisoprodol | 261.05 | 176.10 | 54.90 |
| Chlordiazepoxide | 300.25 | 226.95 | 282.00 |
| Clonazepam | 316.20 | 270.00 | 240.85 |
| Cocaethylene | 318.20 | 195.95 | 81.90 |
| Cocaine | 304.05 | 182.00 | 81.90 |
| Codeine | 300.10 | 215.00 | 165.00 |
| Dextromethorphan | 272.10 | 215.05 | 147.05 |
| Diazepam | 285.25 | 193.05 | 153.95 |
| Dibutylone | 236.10 | 85.95 | 160.90 |
| Dihydrocodeine | 302.10 | 199.00 | 171.00 |
| EDDP | 279.00 | 235.05 | 250.00 |
| Estazolam | 295.10 | 266.95 | 204.95 |
| Etizolam | 343.10 | 313.85 | 258.85 |
| Fentanyl | 337.20 | 188.10 | 104.90 |
| Flunitrazepam | 314.00 | 268.10 | 238.95 |
| Flurazepam | 388.40 | 315.00 | 287.85 |
| Hydrocodone | 300.10 | 199.00 | 128.00 |
| Hydromorphone | 286.05 | 185.00 | 157.00 |
| Ketamine | 328.25 | 124.95 | 220.10 |
| Lorazepam | 322.90 | 276.90 | 304.90 |
| MDA | 180.10 | 162.90 | 104.95 |
| MDEA | 209.00 | 164.00 | 105.90 |
| MDMA | 195.10 | 163.95 | 105.95 |
| MDPV | 277.10 | 206.00 | 175.95 |
| Meperidine | 249.10 | 221.05 | 175.00 |
| Meprobamate | 219.05 | 158.05 | 54.90 |
| Methadone | 310.20 | 265.00 | 104.90 |
| Methamphetamine | 150.05 | 90.95 | 118.95 |
| Methylone | 208.05 | 160.05 | 190.00 |
| Methylphenidate | 234.10 | 83.90 | 56.00 |
| Midazolam | 326.30 | 291.10 | 249.05 |
| Morphine | 286.05 | 152.90 | 164.95 |
| N-Ethylpentalone | 250.10 | 201.95 | 231.95 |
| Nitrazepam | 282.00 | 236.05 | 180.00 |
| N-methyl-norfentanyl | 247.10 | 97.95 | 55.00 |
| Nordiazepam | 271.30 | 139.95 | 207.95 |
| Norhydrocodone | 286.05 | 198.95 | 240.90 |
| Noroxycodone | 301.95 | 283.95 | 227.05 |
| Noroxymorphone | 288.00 | 270.00 | 212.95 |
| O-desmethyl-tramadol | 251.10 | 57.95 | 58.95 |
| Oxazepam | 288.90 | 243.00 | 270.95 |
| Oxycodone | 316.10 | 298.10 | 241.05 |
| Oxymorphone | 302.05 | 283.95 | 226.90 |
| Pentobarbital | 225.00 | 182.05 | 42.00 |
| Phenazepam | 350.90 | 206.00 | 179.00 |
| Phencyclidine | 245.10 | 85.95 | 90.90 |
| Phenobarbital | 231.00 | 42.00 | 188.00 |
| Phentermine | 150.05 | 90.90 | 64.95 |
| Prazepam | 325.40 | 271.1 | 139.95 |
| Secobarbital | 237.00 | 194.00 | 42.00 |
| Tapentadol | 222.15 | 106.95 | 120.95 |
| Temazepam | 301.25 | 255.15 | 282.85 |
| Tramadol | 265.10 | 58.00 | 58.95 |
| Zolpidem | 308.10 | 235.10 | 262.90 |
| 4-ANPP-D5 | 286.25 | 187.95 | - |
| 6-MAM-D6 | 334.10 | 165.00 | - |
| 7-Aminoclonazepam-D4 | 290.00 | 121.00 | - |
| Alpha-PVP-D8 | 241.00 | 90.90 | - |
| Alprazolam-D5 | 314.05 | 285.90 | - |
| Amphetamine-D11 | 147.10 | 98.00 | - |
| Benzoylecgonine-D8 | 298.10 | 170.95 | - |
| Buprenorphine-D4 | 472.15 | 58.95 | - |
| Butylone-D3 | 224.90 | 176.95 | - |
| Carfentanil-D5 | 400.00 | 340.00 | - |
| Carisoprodol-D7 | 268.15 | 183.10 | - |
| Chlordiazepoxide-D5 | 305.00 | 288.00 | - |
| Clonazepam-D4 | 320.00 | 273.95 | - |
| Cocaethylene-D3 | 321.20 | 198.95 | - |
| Cocaine-D3 | 307.00 | 184.95 | - |
| Codeine-D6 | 306.10 | 217.95 | - |
| Dextromethorphan-D3 | 275.15 | 215.00 | - |
| Diazepam-D5 | 290.00 | 198.05 | - |
| Dihydrocodeine-D6 | 308.10 | 202.00 | - |
| EDDP-D3 | 282.10 | 234.95 | - |
| Estazolam-D5 | 300.30 | 272.05 | - |
| Etizolam-D3 | 346.00 | 316.95 | - |
| Fentanyl-D5 | 342.15 | 104.90 | - |
| Flunitrazepam-D7 | 321.05 | 274.95 | - |
| Hydrocodone-D6 | 306.10 | 202.00 | - |
| Hydromorphone-D6 | 292.10 | 184.90 | - |
| Ketamine-D4 | 238.25 | 128.95 | - |
| MDA-D5 | 185.05 | 168.05 | - |
| MDEA-D6 | 214.10 | 166.05 | - |
| MDMA-D5 | 200.05 | 166.05 | - |
| MDPV-D8 | 285.10 | 135.05 | - |
| Meperidine-D4 | 252.10 | 178.00 | - |
| Methadone-D9 | 319.20 | 267.95 | - |
| Methamphetamine-D11 | 161.00 | 96.90 | - |
| Methylone-D3 | 212.00 | 164.05 | - |
| Methylphenidate-D9 | 243.00 | 92.95 | - |
| Midazolam-D4 | 330.00 | 294.95 | - |
| Morphine-D6 | 292.10 | 152.00 | - |
| Nordiazepam-D5 | 276.00 | 139.95 | - |
| Norhydrocodone-D3 | 289.10 | 202.00 | - |
| Noroxycodone-D3 | 306.10 | 287.95 | - |
| O-desmethyl-tramadol-D6 | 256.10 | 64.00 | - |
| Oxazepam-D5 | 292.00 | 245.90 | - |
| Oxycodone-D6 | 322.10 | 303.95 | - |
| Oxymorphone-D6 | 305.05 | 287.05 | - |
| Phenazepam-D4 | 354.90 | 209.95 | - |
| Phencyclidine-D5 | 249.15 | 96.00 | - |
| Phentermine-D5 | 155.15 | 96.00 | - |
| Prazepam-D5 | 330.35 | 275.95 | - |
| Tapentadol-D3 | 225.10 | 106.95 | - |
| Temazepam-D5 | 306.00 | 259.95 | - |
| Tramadol-13C-D3 | 268.15 | 57.90 | - |
| Zolpidem-D6 | 314.10 | 234.90 | - |
| Butalbital-D5 | 228.00 | 42.05 | - |
| Pentobarbital-D5 | 230.00 | 186.95 | - |
| Phenobarbital-D5 | 235.95 | 42.05 | - |
| Secobarbital-D5 | 242.00 | 199.00 | - |

    - The Multi-Drug Screen-Negative method is a screening method for acidic drugs.

    - The Multi-Drug Quant-Negative method is a quantitative method for acidic drugs.

  - Chromatograph and quantitate the calibration standards and QC standards on the LCMS-MS using the appropriate method.

  - Save the method with reference to the specific method used followed by an eight digit format year/month/day.

Examples:  Cannabinoid20190101, MultiDrug20200101

  - Update the calibration with the calibration standards.

  - The response at each level shall be the response of the calibration standard at that level.

  - The calibration curve will be fitted to a linear model.

  - Save the data analysis method.

  - Evaluate the data according to the following acceptance criteria:

    - A calibration point(s) may be dropped from the calibration model if it is determined to be unacceptable provided that 75% of the calibration levels remain.  Points are considered unacceptable for any of the following conditions:

      - A calibration point shows a residual value of >20% from the current calibration line for concentrations greater than or equal to 10 ng/ml and > 30% for concentrations less than 10 ng/ml.

      - Ion ratios are > 20% from the average of the calibration standards.

      - A calibration point response is overloaded to a degree that it affects the upper limit of linearity of the calibration model.

    - The negative calibration standard must not contain a reportable amount of drug targeted by the method used.

      - If the negative calibration standard contains a reportable amount of a drug, that drug cannot be reported for any cases in that extraction batch.

    - The coefficient of determination, r2, must be 0.985 or greater for each component.

  - If the calibration does not meet the acceptance criteria, appropriate action, e.g., maintenance or new solution preparation, shall be taken and the calibration repeated.  If the linearity does not extend over the entire calibration range, then quantitative values may be reported only for the linear range.

- Calibration QC

  - Quantitate the QC standards with the data analysis method from 7.3.  Evaluate according to the acceptance criteria:

    - Drugs to be reported must not be identified at a reportable amount in the negative QC standard.

    - The retention times of each drug and the internal standards must not differ by more than 5.0 % from the target value.

    - For Quantitative method analysis, the qualifier ion ratios of each drug and the internal standards must be within +/- 20 % of the target value.

    - For Quantitative method analysis, the measured quantitation results of each QC sample within the linear range for a component must be within +/- 20 % of the target value for all concentrations above 10 ng/ml.  For concentrations between 1 ng/ml and 10 ng/ml the measured quantitation results of each component must be within +/- 30 % of the target value.

    - If the calibration does not meet the acceptance criteria, appropriate action, e.g., maintenance or new solution preparation, shall be taken and the calibration repeated.

  - Prepare a folder labeled with the name of the calibrated method, refer to 7.3.  Maintain this file in the DWI Blood Chemistry Unit.  This file shall include:

    - Quantitation reports for each calibration and QC standards.

    - Printed calibration table.

    - Each component r2 value and the equation describing each calibration curve.

  - The calibration data and Calibration QC data shall be reviewed by another chemist for acceptance prior to evaluation of case samples.

- Quality Control

  - Internal standards are added to each extraction sample.  The internal standards must be identified in each extraction sample.

  - The negative blood QC must not contain a reportable amount of drugs to be reported.

  - The verification standard must be prepared and run in each batch of cases.

  - All of the positive QC standards and verifier sample(s) must meet the criteria in 8.1.  Analytes  that do not meet the criteria in 8.1.4 may not be reported quantitatively but may be reported as a qualitative identification.

  - Use the current date in the names of batch sequences.

  - Include the CCBI case file number in the data file name and any additional information needed to uniquely identify the sample.

  - Data files associated with casework and performance checks shall not be deleted or overwritten.

  - Notify the DWI Blood Chemistry Technical Leader if the data storage location becomes full.

- Evaluation of Data

  - The LCMS-MS provides retention time data and mass spectral data.

  - For Multiple Reaction Monitoring methods evaluate the chromatogram and spectra for each peak of interest.

  - Positive LCMS-MS Comparison Criteria:

    - The ratios of the quantitation and the qualifier ion(s) must be within +/- 20 % of their expected ratios.

    - The retention times of each drug and the internal standards must not differ by more than 5.0% from the target value.

    - The measured concentration must meet or exceed the lower limit of quantitation (LLOQ) prior to rounding.  See Drug Reporting Criteria for quantitation ranges.

      - Drugs that have measured concentrations below the LLOQ and have positive results with two other tests, will be reported with the statement “(Insert drug identified), below (LLOQ value)” if all other qualitative criteria for identification are met.

  - Record in the case record the:

    - Reference to the QCC file

    - Sample quantitation report

    - Mass spectra of significant peaks of interest, including their internal standard.

- Shutdown / Startup

  - If the LCMS-MS is not going to be used for an extended amount of time (multiple days) the pin shall be inserted into the MS and the Nitrogen generator shall be turned off.

  - The computer may be shut down or restarted if needed.

- Limitations

  - The screening LCMS-MS methods described in this procedure are not confirmatory tests and reported identifications must be confirmed either by GC-MS or LCMS-MS.

  - Analytes are validated on an analyte-specific basis.  See “Drug Reporting Criteria” for all specific reporting criteria for each analyte.

  - Analytes that are not contained in the verification standard cannot have a quantitative value reported.  Analytes identified may be reported qualitatively provided they meet the criteria in 10.3.3.

- Safety

  - Refer to the CCBI Health and Safety Manual.

  - Use extreme caution handling/installing/removing/transporting compressed gas cylinders.  Cylinders shall not be moved without the cylinder cap securely in place.

  - Liquid Chromatograph and Mass Spectrometer may be extremely hot.  Avoid touching hot areas and wear protective gloves while performing maintenance.

- References

  - Nexera X2 SIL-30AC, Shimadzu Corporation, Dec. 2017.

  - Nexera X2 Safety Guideline, Shimadzu Corporation, Jan. 2015.

  - Nexera X2 System Guide, Shimadzu Corporation, June 2013.

  - High Pressure Flow Channel Selection Valve for Shimadzu High Performance Liquid Chromatograph Instruction Manual, Shimadzu Corporation, Sept. 2017.

  - LCMS-8050 Instruction Manual, Shimadzu Corporation, August 2015.

  - HIPACE 300 Turbopump Operating Instructions, Pfeiffer Vacuum.

  - SPILTFLOW 310 Supplementary Information, Pfeiffer Vacuum.

  - NitroFlow60 and NitroFlow60H Installation, Operation, and Maintenance Manual, Parker Balston, 2016.

  - Moffat, A. C., et al., eds. Clarke’s Isolation and Identification of Drugs. 2nd Edition. London:  Pharmaceutical Press, 1986.

  - Moffat, A.C., et al., eds. Clarke’s Analysis of Drugs and Poisons. 4rd Edition. London: Pharmaceutical Press, 2011.

  - Guide for the Use of the International System of Units (SI). NIST Special Publication 811, 2008 Ed., (March 2008; 2nd printing November 2008). p.43

  - Skoog, Douglas A., F. James Holler and Timothy A. Nieman.  Principles of Instrumental Analysis. 5th Edition. Garcourt Brace & Company, 1998.

- Records

  - LCMS-MS Maintenance Log

  - LCMS-MS Activity Log

  - Case record

  - Reagent Log

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 10/28/2019 | 1 | Initial procedure. |
| 01/24/2020 | 2 | 3. Abbreviations - removed abbreviations 4.1 Equipment - added nitrogen generator 4.3 Materials - removed ToxBox Kits 5.5 Monthly QCC- Checktune - removed section 7. Procedures - added 7.2 through 10. from DBCTP11 |
| 5/18/20 | 3 | 7.1.1.5., 7.1.2.5, 7.1.3.5., 7.1.4.5. - Removed collision energy, Q1 and Q3 bias from tables. 7.1.3.4. and 7.1.4.4 MultiDrug Screen and Quant LCMS-MS conditions - Updated Mass spectrometer mode to include negative electrospray ionization for acidic drugs 7.8 Acceptance Criteria - Revised 7.8.1 through 7.8.3 9. Quality Control - Added to 9.4 to allow for analytes to be reporting qualitatively 10.3 Positive LCMS-MS Comparison Criteria - changed limit of detection/limit of quantitation to limit of reporting. 12. Limitations - Added 12.3 to address qualitatively identifying analytes which are identified but not contained in the verification standard. |
| 10/26/20 | 4 | 10.3 Identification of Drugs – changed title to Evaluation of Data, changed limit of reporting to LLOQ in 10.3.3. Added 10.3.3.1. to allow for reporting of substances below LLOQ when qualitative criteria is met, and two other positive tests have been performed. |

# 16: Technical Procedure for Supported Liquid Extraction (SLE) of Acidic, Basic, and Neutral Drugs for LCMS-MS Analysis

- Purpose / Scope - This procedure is used to isolate acidic, basic, and neutral drugs from blood using a ToxBox custom analytical plate and a Thermo 96 well SLE plate for analysis by Liquid chromatography - mass spectrometry in the DWI Blood Chemistry Unit of the CCBI Crime Laboratory.

- Definitions

  - Quality control check (QCC) - Periodic confirmation of the reliability of equipment, instrumentation, and/or reagents.

- Abbreviations

  - Refer to DWI Blood Chemistry Unit Technical Procedure for Analysis

- Equipment, Materials and Reagents

  - Equipment

    - Biotage SPE Dry Evaporator

    - Positive Pressure Manifold

    - Shaking Incubator

    - Mechanical pipettors and corresponding tips

    - Graduated Cylinder

  - Materials

    - Vortex mixer

    - Thermo 96 well Supported Liquid Extraction Plate

    - Deionized water

    - Nitrogen, UHP

  - Reagents

    - Critical Reagents

      - Negative blood

    - Reference materials

      - ToxBox MultiDrug custom analytical plate containing selected acidic, basic, and neutral drugs.

    - Commercial reagents

      - Formic Acid, LC/MS grade

      - Methanol, LC/MS grade

      - Ethyl Acetate, LC/MS grade

      - Ammonium Hydroxide, ACS grade

      - Ammonium Formate, ACS grade

    - Prepared Reagents

      - 0.5M Ammonium Hydroxide

        - Mix 466.4 ml of deionized water and 33.6 ml of Ammonium Hydroxide.

        - Lot Number:  Eight digit format year/month/day/0.5MAmmoniumHydroxide /initials of preparer.

        - Expiration:  1 month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - MultiDrug Screening LCMS-MS Mobile Phase A - 0.1% Formic Acid in Water

        - Mix Formic Acid and deionized water in a ratio of 1000:1.

        - Lot Number:  Eight digit format year/month/day/MPA0.1FAH2O/initials of preparer.

        - Expiration:  1 month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - MultiDrug Screening LCMS-MS Mobile Phase B - 0.1% Formic Acid in Methanol

        - Mix Methanol and Formic acid in a ratio of 1000:1.

        - Lot Number:  Eight digit format year/month/day/MPB0.1%FAMeOH/initials of preparer.

        - Expiration:  one month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - Quantitative LCMS-MS Mobile Phase A - 10 mM Ammonium Formate in Water

        - Mix 0.631 gram of Ammonium Formate in 1000 ml of deionized water.

        - Lot Number:  Eight digit format year/month/day/MPA10mMAmmFormateH2O/initials of preparer.

        - Expiration:  1 month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - Quantitative LCMS-MS Mobile Phase B - 0.1% Formic Acid in Methanol

        - Mix Methanol and Formic acid in a ratio of 1000:1.

        - Lot Number:  Eight digit format year/month/day/MPB0.1%FAMeOH/initials of preparer.

        - Expiration:  one month

        - Storage:   closed container

        - QCC:  Successful QC positive control

      - Multi-Drug Verification Solution

        - Prepare this solution with reference materials from a different supplier or different supplier lot than those used in the ToxBox MultiDrug custom analytical plates.

        - Prepare a standard solution that contains components and concentrations of reference material as specified in the Verification Solution Log located on the shared drive.

        - Lot Number: Eight digit format year/month/day/MultiDrugVerStd/unique identifier(if applicable)/initials of preparer.

        - Expiration: 	one year

        - Storage: 	refrigerator

      - Multi-Drug Verifier

        - Prepare a verifier that contains components and concentrations of reference material as specified in the Verification Solution Log located on the shared drive.

        - Lot Number: Eight digit format year/month/day/MultiDrugVer/unique identifier (if applicable)/initials of preparer.

        - Expiration:	end of the day

        - Storage:	closed container

        - QCC: Successful QC positive control

- Calibration

  - Calibrate the instrument each time that an extraction batch is analyzed and after instrument maintenance that may affect the calibration as determined by the DWI Blood Chemistry Technical Leader.

  - Calibration Standards

    - Screening plate for Basic Drugs (standard dosage):

Level 1 - 5 ng/ml of each drug

Level 2 - 25 ng/ml of each drug

Level 3 - 100 ng/ml of each drug

Level 4 - 250 ng/ml of each drug

The QC levels are the following:

QC1 - 10 ng/ml of each drug

QC2 - 25 ng/ml of each drug

The internal standard concentration is 50 ng/ml of each drug (deuterated).

    - Screening plate for Basic Drugs (low dosage):

Level 1 - 0.5 ng/ml of each drug

Level 2 - 2.5 ng/ml of each drug

Level 3 - 10 ng/ml of each drug

Level 4 - 25 ng/ml of each drug

The QC levels are the following:

QC1 - 1 ng/ml of each drug

QC2 - 2.5 ng/ml of each drug

The internal standard concentration is 50 ng/ml of each drug (deuterated).

    - Screening plate for Barbiturates:

Level 1 - 50 ng/ml of each drug

Level 2 - 250 ng/ml of each drug

Level 3 - 1000 ng/ml of each drug

Level 4 - 2500 ng/ml of each drug

The QC levels are the following:

QC1 - 100 ng/ml of each drug

QC2 - 250 ng/ml of each drug

The internal standard concentration is 500 ng/ml of each drug (deuterated).

    - Quantitation plate for Basic Drugs (standard dosage):

Level 1 - 5 ng/ml of each drug

Level 2 - 10 ng/ml of each drug

Level 3 - 25 ng/ml of each drug

Level 4 - 50 ng/ml of each drug

Level 5 - 100 ng/ml of each drug

Level 6 - 250 ng/ml of each drug

Level 7 - 500 ng/ml of each drug

Level 8 - 1000 ng/ml of each drug

The QC levels are the following:

QC Low - 10 ng/ml of each drug

QC Mid 1 - 100 ng/ml of each drug

QC Mid 2 - 250 ng/ml of each drug

QC High - 1000 ng/ml of each drug

The internal standard concentration is 50 ng/ml of each drug (deuterated).

    - Quantitation plate for Basic Drugs (low dosage):

Level 1 - 0.5 ng/ml of each drug

Level 2 - 1 ng/ml of each drug

Level 3 - 2.5 ng/ml of each drug

Level 4 - 5 ng/ml of each drug

Level 5 - 10 ng/ml of each drug

Level 6 - 25 ng/ml of each drug

Level 7 - 50 ng/ml of each drug

Level 8 - 100 ng/ml of each drug

The QC levels are the following:

QC Low - 1 ng/ml of each drug

QC Mid 1 - 10 ng/ml of each drug

QC Mid 2 - 25 ng/ml of each drug

QC High - 100 ng/ml of each drug

The internal standard concentration is 50 ng/ml of each drug (deuterated).

    - Quantitation plate for Barbiturates:

Level 1 - 50 ng/ml of each drug

Level 2 - 100 ng/ml of each drug

Level 3 - 250 ng/ml of each drug

Level 4 - 500 ng/ml of each drug

Level 5 - 1,000 ng/ml of each drug

Level 6 - 2,500 ng/ml of each drug

Level 7 - 5,000 ng/ml of each drug

Level 8 - 10,000 ng/ml of each drug

The QC levels are the following:

QC Low - 100 ng/ml of each drug

QC Mid 1 - 1,000 ng/ml of each drug

QC Mid 2 - 2,500 ng/ml of each drug

QC High - 10,000 ng/ml of each drug

The internal standard concentration is 500 ng/ml of each drug (deuterated).

- Procedure

  - Remove analytical plate, blood and samples from cold storage and allow to reach room temperature.

  - Ensure that all blood samples are homogenous by shaking and/or vortexing.

    - If a homogenous sample cannot be obtained, make a notation in the worksheet detailing the condition of the sample and its handling.

  - Add 0.25 mL negative blood to wells of analytical plate containing standards/QCs; add 0.25 mL of case sample blood to wells containing only internal standard.  Ensure that the blood samples are added consistent with the plate map.

  - Place plate on the shaking incubator at approximately 900 rpm for 15 minutes.

  - Add 0.25 mL 0.5 M Ammonium Hydroxide to all wells.

  - Place plate on the shaking incubator at approximately 900 rpm for 15 minutes.

  - Transfer 300 µL of the blood + Ammonium Hydroxide mixture to corresponding wells of a Thermo SLE plate.

  - Apply positive pressure until solution penetrates well sorbent approximately 50% (or until no liquid remains on top of sorbent) – approximately 4 seconds.

  - Wait 5 minutes for the sample to completely absorb.

  - Add 0.900 mL Ethyl Acetate to all wells and allow to flow for 5 minutes under gravity.

  - Apply positive pressure to complete elution – approximately 20 seconds.

  - Add 0.900 mL Ethyl Acetate and allow to flow for 5 minutes under gravity.

  - Apply positive pressure to complete elution – approximately 20 seconds.

  - Evaporate to dryness in the SPE Dry at approximately 35°C under a constant flow of nitrogen.  Temperatures for this step are not critical and it is appropriate to use equipment temperature readings.

  - Reconstitute samples and heat-seal foil on top to prevent evaporation.

    - For MultiDrug screen analysis, reconstitute each sample with 100 µL of 20 % Methanol in deionized water.

    - For MultiDrug quant analysis, reconstitute each sample with 100 µL of Methanol.

  - Analyze all samples immediately or refrigerate for future analysis.

- Limitations

  - Extracted plates are stable for 7 days from the extraction date.

- Safety

  - Refer to the CCBI Health and Safety Manual

- References

  - Biotage Pressure+ 96 User Manual, Biotage, Uppsala, Sweden.SPE Dry SPE Dry 96 Dual User Manual, Bke Touch Instruction Manual, Fisher Scientific, Revision 8, 2/2018.

- Records

  - Multi Drug QCC file

  - Reagent Log

  - Verification Solution Log

| Revision History |  |  |
| --- | --- | --- |
| Effective Date | Version Number | Reason |
| 05/18/2020 | 1 | New Procedure |